## 480 (Jan '07)
Transmissometer Calibration
Serial Number: |  | CST-480 DR |  | Factory Calibration Date: |  | 2007-01-05
Calibration coefficients M and B are entered into the Sea-Bird CTD instrument configuration (.con) file. They are calculated, as per SBE Application Note 7 (March 2004), as follows:
M = (Tw / [W0 - Y0]) * (A0 - Y0) / (A1 - Y1)
B = -1 * M * Y1
Where:
A0 = | 4.899 | (Air calibration voltage from WetLabs calibration sheet)
Y0 = | 0.059 | (Blocked path voltage from WetLabs calibration sheet)
Tw = | 90.2 | (Percentage transmission in pure water - See SBE app. note 7, March 2004)
W0 = | 4.788 | (Voltage output in pure water from WetLabs calibration sheet)
A1 = | VAir | (Current air voltage)
Y1 = | VBlocked | (Current blocked path voltage)
Field Recalibration
Date | VAir | VBlocked | M | B | Notes
2007-01-05 | 4.899 | 0.059 | 19.0738 | -1.1254 | Factory values 05-Jan-07
2007-01-13 | 4.7655 | 0.0583 | 19.6119 | -1.1434 | LAB - Beg EN427 427_xaooF1.*
2007-02-20 | 4.951 | 0.0586 | 18.8695 | -1.1058 | WLF - Begin EN429, Ullman
2007-04-24 | 4.8339 | 0.0574 | 19.3274 | -1.1094 | WLF - Begin EN433, Martin
2007-05-16 | 4.735 | 0.0586 | 19.7411 | -1.1568 | LAB - Beg EN435 435xapooFs1.*
2007-05-31 | 4.8095 | 0.0586 | 19.4315 | -1.1387 | LAB - End EN435 435xapooFs6.*

## 480 (Feb '06)
Transmissometer Calibration
Serial Number: |  | CST-480 DR |  | Factory Calibration Date: |  | 2006-02-24
Calibration coefficients M and B are entered into the Sea-Bird CTD instrument configuration (.con) file. They are calculated, as per SBE Application Note 7 (March 2004), as follows:
M = (Tw / [W0 - Y0]) * (A0 - Y0) / (A1 - Y1)
B = -1 * M * Y1
Where:
A0 = | 4.882 | (Air calibration voltage from WetLabs calibration sheet)
Y0 = | 0.057 | (Blocked path voltage from WetLabs calibration sheet)
Tw = | 90.2 | (Percentage transmission in pure water - See SBE app. note 7, March 2004)
W0 = | 4.782 | (Voltage output in pure water from WetLabs calibration sheet)
A1 = | VAir | (Current air voltage)
Y1 = | VBlocked | (Current blocked path voltage)
Field Recalibration
Date | VAir | VBlocked | M | B | Notes
2006-02-24 | 4.867 | 0.062 | 19.1694 | -1.1885 | Factory values 24-Feb-06
2006-04-08 | 4.7192 | 0.0586 | 19.7633 | -1.1581 | LAB - Beg EN416 416_xFooaps1.*
2006-04-15 | 4.7888 | 0.0586 | 19.4725 | -1.1411 | LAB - End EN416 416_xFoos2.*

## 480 (Dec '01)
Transmissometer Calibration
Serial Number: |  | CST-480 DR |  | Factory Calibration Date: |  | 2001-12-21
Calibration coefficients M and B are entered into the Sea-Bird CTD instrument configuration (.con) file. They are calculated, as per SBE Application Note 7 (March 2004), as follows:
M = (Tw / [W0 - Y0]) * (A0 - Y0) / (A1 - Y1)
B = -1 * M * Y1
Where:
A0 = | 4.882 | (Air calibration voltage from WetLabs calibration sheet)
Y0 = | 0.057 | (Blocked path voltage from WetLabs calibration sheet)
Tw = | 90.2 | (Percentage transmission in pure water - See SBE app. note 7, March 2004)
W0 = | 4.782 | (Voltage output in pure water from WetLabs calibration sheet)
A1 = | VAir | (Current air voltage)
Y1 = | VBlocked | (Current blocked path voltage)
Field Recalibration
Date | VAir | VBlocked | M | B | Notes
2001-12-21 | 4.882 | 0.057 | 19.0899 | -1.0881 | Factory values
2002-02-11 | 4.8632 | 0.0562 | 19.1614 | -1.0769 | LAB - Mid EN366
2002-03-21 | 4.8767 | 0.0562 | 19.1078 | -1.0739 | LAB - End EN367
2002-08-20 | 4.8278 | 0.0562 | 19.3036 | -1.0849 | LAB - Post EN374
2002-09-20 | 4.8071 | 0.0562 | 19.3877 | -1.0896 | LAB - Start EN376
2003-03-17 | 4.7949 | 0.0574 | 19.4425 | -1.116 | LAB - EN378
2003-05-23 | 4.7961 | 0.0562 | 19.4327 | -1.0921 | LAB - EN380b
2003-07-08 | 4.7581 | 0.0562 | 19.5897 | -1.1009 | LAB - End EN383
2003-08-11 | 4.7595 | 0.0562 | 19.5839 | -1.1006 | LAB - Start EN385
2003-08-18 | 4.7668 | 0.0562 | 19.5536 | -1.0989 | LAB - End EN385
2005-03-08 | 4.7167 | 0.0598 | 19.779 | -1.1828 | LAB - Start EN402 402xfaoo4.* Used for EN403
2005-03-15 | 4.6778 | 0.0574 | 19.9353 | -1.1443 | LAB - End EN402 402xfFoo6.*
2005-03-31 | 4.7301 | 0.0608 | 19.7265 | -1.1994 | LAB - Mid EN403 403xfaoos4.*
2005-04-05 | 4.7164 | 0.0647 | 19.8011 | -1.2811 | LAB - End EN403 403xfapoos5.*