CTD SYSTEM CONFIGURATION			Cruise:	EN448	Dates:	27 Jun-03Jul, 2007

Tech: Lynne Butler					Chief Scientist, Institution: McGillicuddy, Anderson, WHOI

Program, versions					SeaSave Win32 v 5.37d, DOS Seasoft 4.249,

# SBE Processing v5.32a

File names						EN448_0607.cfg  CTD001-CTD067.*

Batch files used	448pro.bat & 448wpro_a.txt to process and archive

Deck cal files		EN448\CTD\sswin\CalData\EN448\  448xapooFs12_0607.dat & 448xapooFs13_0607.dat

SBE 9plus CTD underwater unit (6800m depth capability)	S/N	12293-0444 (07-Mar-03), 23539-0607 (03-Mar-06), none

Pressure sensor, Digiquartz (with temp comp, SBE)		S/N	64853 (07-Mar-03), 80016 (03-Mar-06)

SBE 11plus deck unit					S/N	11P12293- 0426, 0427 (21-Mar-97), none

Temperature sensor, primary (3400 m, 6800m, slow, fast)	S/N	604, 753, 1164, 1185, 1678, 2034, 2107,

(SBE 3, 3Plus)				2899, 2902, 4126, 4130, 4333, 4695

Temperature sensor, secondary (3400 m, 6800m, slow, fast)	S/N	604, 753, 1164, 1185, 1678, 2034, 2107,

(SBE 3, 3Plus)				2899, 2902, 4126, 4130, 4333, 4695, none

Conductivity sensor, primary (SBE 4C, 3400m, 6800m)	S/N	199, 200, 618, 864, 1446, 1578

1749, 2459, 2469, 2822, 3220

Conductivity sensor, secondary (SBE 4C, 3400m, 6800m)	S/N	199, 200, 618, 864, 1446, 1578,

1749, 2459, 2469, 2822, 3220, none

Pump, primary	(SBE 5T, 10500m, 3000RPM)		S/N	1413, 1595, 1596, 2238,

2825, 2892, 3566, 3572

Pump, secondary	(SBE 5T, 10500m, 3000RPM)		S/N	1413, 1595, 1596, 2238,

2825, 2892, 3566, 3572, none

Oxygen sensor, primary (SBE 13, SBE 23, SBE 43, 6800m)	S/N	430343, 430345, 430348, 430372,

431230, 431339, none

Oxygen sensor, secondary (SBE 13, SBE 23, SBE 43, 6800m)	S/N	430343, 430345, 430348, 430372,

431230, 431339, none

Fluorometer, primary	(Sea Tech flashlamp, 3000m,	S/N	117S (30Nov04 10microg/l full scale), V6=4.923

WET Labs ECOFluor, 6000m)		FLRTD-231, 478, 492, none

Fluorometer, secondary	(Sea Tech flashlamp, 3000m)	S/N	117S, FLRTD-231, 478, 492, none

WET Labs ECOFluor, 6000m)

Transmissometer (WET Labs C-Star, 25 cm path, 6000m)	S/N	399DR, 480DR, 593DR, 969DR, none

Altimeter (Benthos PSA-916, 2500m)  deck V2=4.926	S/N	1017, 1075, none

Light sensor PAR, QSP-200L4S/2300 (Biospherical, 1000m)	S/N	4479, 4701, none

Light sensor, Surface SPAR, QSR-2200 (Biospherical)	S/N	20121, 20190, none

Data format (factory & user installed sensors) See also EN448\CTD\CTDReadMe448.txt

Frequency 0		temperature, primary 4695		External voltage 3 	PAR 4701

Frequency 1		conductivity, primary 2469	External voltage 4 	DO1 1230

Frequency 2		pressure		80016		External voltage 5 	DO2 0345

Frequency 3		temperature, secondary 2902	External voltage 6	spare

Frequency 4		conductivity, secondary 1749	External voltage 7	Fluor 231

External voltage 0 4.7488	Transmissometer 399DR		External voltage 8	spare

External voltage 1	spare				External voltage 9 	SPAR 20121

External voltage 2 4.926	Altimeter 1075			Pump1=      , Pump2=      .

CTD SYSTEM CONFIGURATION, Continued

Carousel	 pylon, rosette, levers		GO, SBE	S/N	3210573-0048, 3216020-0175, shiny12, dull24, none

Frame w/ s/s cage							12 pl SBE alum old/new, 12 pl s/s, 24pl SBE alum, bk1, bk2

Niskins		# on frame: 12 	      1.5, 5, 10, 30 L	Go-Flo, Ext resist OTE / GO, Internal resist, w/ Hiflow Valves, none

Winch / wire						#	1in, 1out, 2, 3, Scanfish / .322”, .500”, .680”, 3/8”, ½”, 9/16”

Slip Rings (Meridian, Focal)				S/N	6.859, 8.538 (10wire), 8.594, Focal 1801308, none

Pinger	(Williams)					S/N	W1, W2, W3, none

(Benthos 2216, 12000m)				S/N	1237, none

Note for SBE 9plus CTD underwater unit S/N 12293-0444, 0523 & 0607:

CTDs 0444, 0523, and 0607 were supplied with an isolated 12-volt power source for a Sea Tech fluorometer in a 3000-meter stainless steel housing. When this fluorometer is used it must be installed on connector JT2 (see drawing 50077). This will put the fluorometer output on A/D channel 1 (External Voltage 1).

Underwater unit S/N 9P16940-0607, as of x x 2000:

Pressure sensor range						0-10,000 psia, 0-6885 d-Bar

pressure sensor, Digiquartz (with temp comp, SBE)		S/N	80016

Modulo							S/N	MOD12P-439

AD590M (enter in Seacon)					0.012837, , .0128452

AD590B (enter in Seacon)						-9.963864, -8.66958

A/D Input Voltage Range						0 to 5 volts DC

Logic Board EPROM				Version 1.0

Modem Interface NOT installed

NO Modem Board Microcontroller			Version 2.0A, Version 1.1		P/N N/A, 11124, 11194

Isolated Fluorometer Power Interface NOT Installed

## Voltages in Deck test439xapooFs1_0607.dat		Open		Blocked		Capped

External voltage 0	Transmissometer 480DR		4.8083		0.0574

External voltage 1	spare

External voltage 2	spare

External voltage 3	spare

External voltage 4	spare

External voltage 5	spare

External voltage 6	spare

External voltage 7	spare

External voltage 8	spare

External voltage 9	no instrument attached

## Voltages in Deck test438xapooFs1_0607.dat		Open		Blocked		Capped

External voltage 0	Transmissometer 480DR		4.8083		0.0574

External voltage 1	spare				0.000

External voltage 2	Alt 1075				4.926

External voltage 3	PAR 4479			~4.1612daylight	0.4615		~1.9389 w/hole in top

External voltage 4	DO 1230			1.355

External voltage 5	DO 0497			1.402

External voltage 6	spare				4.923

External voltage 7	Fluor 231			0.0611		4.9145w/fluor stick  1.538

External voltage 8	spare

External voltage 9	SPAR 20190			4.497daylight	4.994