## 399 (Apr '08)
Transmissometer Calibration
Serial Number: |  | CST-399 DR |  | Factory Calibration Date: |  | 2008-04-25
Calibration coefficients M and B are entered into the Sea-Bird CTD instrument configuration (.con) file. They are calculated, as per SBE Application Note 7 (March 2004), as follows:
M = (Tw / [W0 - Y0]) * (A0 - Y0) / (A1 - Y1)
B = -1 * M * Y1
Where:
A0 = | 4.8 | (Air calibration voltage from WetLabs calibration sheet=Vair WET Labs)
Y0 = | 0.061 | (Blocked path voltage from WetLabs calibration sheet=Vd WET Labs)
Tw = | 90.2 | (Percentage transmission in pure water - See SBE app. note 7, March 2004)
W0 = | 4.693 | (Voltage output in pure water from WetLabs calibration sheet=Vref WET Labs)
A1 = | VAir | (Current air voltage)
Y1 = | VBlocked | (Current blocked path voltage)
Field Recalibration
Date | VAir | VBlocked | M | B | Notes
2008-04-25 | 4.8 | 0.061 | 19.4732 | -1.1879 | WLF: Calibration values
2008-04-29 | 4.778 | 0.057 | 19.5475 | -1.1142 | DAN: Pre EN-445
2008-06-27 | 4.7448 | 0.0574 | 19.6876 | -1.1301 | LAB: Beg EN448 McGillicuddy

## 399 (Jan '07)
Transmissometer Calibration
Serial Number: |  | CST-399 DR |  | Factory Calibration Date: |  | 2007-01-04
Calibration coefficients M and B are entered into the Sea-Bird CTD instrument configuration (.con) file. They are calculated, as per SBE Application Note 7 (March 2004), as follows:
M = (Tw / [W0 - Y0]) * (A0 - Y0) / (A1 - Y1)
B = -1 * M * Y1
Where:
A0 = | 4.792 | (Air calibration voltage from WetLabs calibration sheet)
Y0 = | 0.058 | (Blocked path voltage from WetLabs calibration sheet)
Tw = | 90.2 | (Percentage transmission in pure water - See SBE app. note 7, March 2004)
W0 = | 4.691 | (Voltage output in pure water from WetLabs calibration sheet)
A1 = | VAir | (Current air voltage)
Y1 = | VBlocked | (Current blocked path voltage)
Field Recalibration
Date | VAir | VBlocked | M | B | Notes
2007-09-30 | 4.695 | 0.057 | 19.872 | -1.1327 | Marshall 30Sep07 In their own CTD (WHOI's)

## 399 g(Feb '05)
Transmissometer Calibration
Serial Number: |  | CST-399 DR |  | Factory Calibration Date: |  | 2005-02-11
Calibration coefficients M and B are entered into the Sea-Bird CTD instrument configuration (.con) file. They are calculated, as per SBE Application Note 7 (March 2004), as follows:
M = (Tw / [W0 - Y0]) * (A0 - Y0) / (A1 - Y1)
B = -1 * M * Y1
Where:
A0 = | 4.877 | (Air calibration voltage from WetLabs calibration sheet)
Y0 = | 0.062 | (Blocked path voltage from WetLabs calibration sheet)
Tw = | 90.2 | (Percentage transmission in pure water - See SBE app. note 7, March 2004)
W0 = | 4.775 | (Voltage output in pure water from WetLabs calibration sheet)
A1 = | VAir | (Current air voltage)
Y1 = | VBlocked | (Current blocked path voltage)
Field Recalibration
Date | VAir | VBlocked | M | B | Notes
2005-02-11 | 4.877 | 0.062 | 19.1386 | -1.1866 | WLF - Factory calibration values
2005-07-26 | 4.8889 | 0.0672 | 19.112 | -1.2843 | LAB - End EN408 Boyle
2005-08-15 | 4.7692 | 0.0562 | 19.5528 | -1.0989 | LAB - Pre EN409 Moran
2005-08-17 | 4.764 | 0.056 | 19.5735 | -1.0961 | WLF - Start EN409 Moran
2005-08-29 | 4.768 | 0.056 | 19.5569 | -1.0952 | WLF - End EN409
2005-09-06 | 4.7619 | 0.0562 | 19.5831 | -1.1006 | LAB - Pre EN410 Kincaid
2005-09-20 | 4.7643 | 0.0562 | 19.5731 | -1.1 | LAB - Pre EN411 Rossby
2005-10-04 | 4.7607 | 0.0549 | 19.5827 | -1.0751 | LAB - Pre EN412 412Fxaoos1.CON

## 399 (Oct '00)
Transmissometer Calibration
Serial Number: |  | CST-399 DR |  | Factory Calibration Date: |  | 2000-10-12
Calibration coefficients M and B are entered into the Sea-Bird CTD instrument configuration (.con) file. They are calculated, as per SBE Application Note 7 (March 2004), as follows:
M = (Tw / [W0 - Y0]) * (A0 - Y0) / (A1 - Y1)
B = -1 * M * Y1
Where:
A0 = | 4.851 | (Air calibration voltage from WetLabs calibration sheet)
Y0 = | 0.058 | (Blocked path voltage from WetLabs calibration sheet)
Tw = | 90.2 | (Percentage transmission in pure water - See SBE app. note 7, March 2004)
W0 = | 4.743 | (Voltage output in pure water from WetLabs calibration sheet)
A1 = | VAir | (Current air voltage)
Y1 = | VBlocked | (Current blocked path voltage)
Field Recalibration
Date | VAir | VBlocked | M | B | Notes
2001-06-05 | 4.8889 | 0.0672 | 19.1383 | -1.2861 | LAB - Mid EN355
2001-07-19 | 4.8413 | 0.0574 | 19.2896 | -1.1072 | LAB - Start EN358
2001-09-16 | 4.8816 | 0.0562 | 19.1237 | -1.0747 | LAB - Start EN360
2001-11-12 | 4.812 | 0.0574 | 19.4084 | -1.114 | LAB - Mid EN363
2004-05-12 | 4.7436 | 0.0562 | 19.6867 | -1.1064 | LAB - Start EN393
2004-11-30 | 4.7387 | 0.0574 | 19.7123 | -1.1315 | LAB - Start EN400