CTD SYSTEM CONFIGURATION


Cruise:
EN451
Dates:
07 – 13 Aug, 2008
Tech: Lynne Butler




Chief Scientist, Institution: McGillicuddy, Anderson, WHOI
Program, versions




SeaSave Win32 v 5.37d, DOS Seasoft 4.249,
SBE Processing v5.32a

File names





EN451_0607.cfg  CTD001-CTD067.*
Batch files used
451pro.bat & 451wpro_a.txt to process and archive

Deck cal files





EN451\CTD\sswin\CalData\EN451\  451xapooFs01_0607_12pl.dat
SBE 9plus CTD underwater unit (6800m depth capability)
S/N
12293-0444 (07-Mar-03), 23539-0607 (03-Mar-06), none

Pressure sensor, Digiquartz (with temp comp, SBE)

S/N
64853 (07-Mar-03), 80016 (03-Mar-06)
SBE 11plus deck unit




S/N
11P12293- 0426, 0427 (21-Mar-97), none

Temperature sensor, primary (3400 m, 6800m, slow, fast)
S/N
604, 753, 1164, 1185, 1678, 2034, 2107,

(SBE 3, 3Plus)



2899, 2902, 4126, 4130, 4333, 4695
Temperature sensor, secondary (3400 m, 6800m, slow, fast)
S/N
604, 753, 1164, 1185, 1678, 2034, 2107,




(SBE 3, 3Plus)



2899, 2902, 4126, 4130, 4333, 4695, none

Conductivity sensor, primary (SBE 4C, 3400m, 6800m)
S/N
199, 200, 618, 864, 1446, 1578









1749, 2459, 2469, 2822, 3220

Conductivity sensor, secondary (SBE 4C, 3400m, 6800m)
S/N
199, 200, 618, 864, 1446, 1578,









1749, 2459, 2469, 2822, 3220, none

Pump, primary
(SBE 5T, 10500m, 3000RPM)

S/N
1413, 1595, 1596, 2238, 









2825, 2892, 3566, 3572

Pump, secondary
(SBE 5T, 10500m, 3000RPM)

S/N
1413, 1595, 1596, 2238, 









2825, 2892, 3566, 3572, none

Oxygen sensor, primary (SBE 13, SBE 23, SBE 43, 6800m)
S/N
430343, 430345, 430348, 430372, 



431230, 431339, none

Oxygen sensor, secondary (SBE 13, SBE 23, SBE 43, 6800m)
S/N
430343, 430345, 430348, 430372, 



431230, 431339, none

Fluorometer, primary
(Sea Tech flashlamp, 3000m,
S/N
117S (30Nov04 10microg/l full scale), V6=4.923




WET Labs ECOFluor, 6000m)

FLRTD-231, 478, 492, none

Fluorometer, secondary
(Sea Tech flashlamp, 3000m)
S/N
117S, FLRTD-231, 478, 492, none
WET Labs ECOFluor, 6000m)

Transmissometer (WET Labs C-Star, 25 cm path, 6000m)
S/N
399DR, 480DR, 593DR, 969DR, none

Altimeter (Benthos PSA-916, 2500m)  deck V2=4.926
S/N
1017, 1075, none

Light sensor PAR, QSP-200L4S/2300 (Biospherical, 1000m)
S/N
4479, 4701, none

Light sensor, Surface SPAR, QSR-2200 (Biospherical)
S/N
20121, 20190, none

Data format (factory & user installed sensors) See also EN451\CTD\CTDReadMe451.txt
Frequency 0

temperature, primary 4695

External voltage 3 
PAR 4701

Frequency 1

conductivity, primary 2469
External voltage 4 
DO1 1230

Frequency 2

pressure

80016

External voltage 5 
DO2 0345

Frequency 3

temperature, secondary 2902
External voltage 6
spare

Frequency 4

conductivity, secondary 1749
External voltage 7
Fluor 231

External voltage 0 4.7534
Transmissometer 399DR

External voltage 8
spare

External voltage 1
spare



External voltage 9 
SPAR 20121

External voltage 2 4.926
Altimeter 1075


Pump1= 2892    Pump2= 2825

CTD SYSTEM CONFIGURATION, Continued

Carousel
 pylon, rosette, levers

GO, SBE
S/N
3210573-0048, 3216020-0175, shiny12, dull24, none

Frame w/ s/s cage






12 pl SBE alum old/new, 12 pl s/s, 24pl SBE alum, bk1, bk2

Niskins

# on frame: 12
1.5, 5, 10, 30 L
Go-Flo, 3,0 Ext resist OTE / GO, 9,12 Internal resist, w/ Hiflow Valves, none

Winch / wire





#
1in, 1out, 2, 3, Scanfish / .322”, .500”, .680”, 3/8”, ½”, 9/16”

Slip Rings (Meridian, Focal)



S/N
6.859, 8.538 (10wire), 8.594, Focal 1801308, none

Pinger
(Williams)




S/N
W1, W2, W3, none
(Benthos 2216, 12000m)



S/N
1237, none
Note for SBE 9plus CTD underwater unit S/N 12293-0444, 0523 & 0607:

  CTDs 0444, 0523, and 0607 were supplied with an isolated 12-volt power source for a Sea Tech fluorometer in a 3000-meter stainless steel housing. When this fluorometer is used it must be installed on connector JT2 (see drawing 50077). This will put the fluorometer output on A/D channel 1 (External Voltage 1).

Underwater unit S/N 9P16940-0607, as of x x 2000:

Pressure sensor range





0-10,000 psia, 0-6885 d-Bar

pressure sensor, Digiquartz (with temp comp, SBE)

S/N
80016

Modulo






S/N
MOD12P-439

AD590M (enter in Seacon)




0.012837, , .0128452

AD590B (enter in Seacon)





-9.963864, -8.66958

A/D Input Voltage Range





0 to 5 volts DC

Logic Board EPROM



Version 1.0

Modem Interface NOT installed

NO Modem Board Microcontroller


Version 2.0A, Version 1.1

P/N N/A, 11124, 11194

Isolated Fluorometer Power Interface NOT Installed

Voltages in Deck test439xapooFs1_0607.dat

Open

Blocked

Capped

External voltage 0
Transmissometer 480DR

4.8083

0.0574

External voltage 1
spare

External voltage 2
spare

External voltage 3
spare

External voltage 4
spare

External voltage 5
spare

External voltage 6
spare

External voltage 7
spare

External voltage 8
spare

External voltage 9
no instrument attached

Voltages in Deck test438xapooFs1_0607.dat

Open

Blocked

Capped

External voltage 0
Transmissometer 480DR

4.8083

0.0574

External voltage 1
spare



0.000

External voltage 2
Alt 1075



4.926

External voltage 3
PAR 4479


~4.1612daylight
0.4615

~1.9389 w/hole in top

External voltage 4
DO 1230


1.355

External voltage 5
DO 0497


1.402

External voltage 6
spare



4.923

External voltage 7
Fluor 231


0.0611

4.9145w/fluor stick  1.538

External voltage 8
spare

External voltage 9
SPAR 20190


4.497daylight
4.994

CTDSysCon451_0607_12pl.doc

17-Mar-08
Page 1

