## If a salinity bottle is right side up it has a sample in it.

## If it is upside down it has already been analyzed.

## Please notify the ship's technician if there is a question about the bottles.

## Instrument power on for 3 hrs w/ bath at room temp

# Fxn: Stdby

## Flush w/TritonX-100

## Flush w/DI

## Ref cal

# FXN: STDBY

REF

Wait 8 sec

Do the following for 10 cycles:

Log the -REFERENCE xxxxx (updates for 16 sec) must be betw 19750-19999   19919, 19918, 19918, 19918, 19919,

Log the +REFERENCE xxxxx (updates for  7 sec) must be betw 19750-19999	19928, 19919, 19918, 19919, 19919

Log the   REFERENCE xxxxx (appears for 8 sec)	39737, 39738, 39738, 39738, 39738

COND		1.297383

## Zero cal

# FXN: ZERO let it settle ~10-15 sec

# COND		-0.000008 –0.000000

ZERO		-0.00023	-0.00023

COND		-0.000000	-0.000000

# FXN: STDBY

## Flush w/Seawater 5-7 times (less than 1 entire salt bottle)

## Standardization

Obtain good cond measurement 3 times:

FXN: STDBY

Rock sample bottle gently

Wipe tube, Install standard sample bottle

Flow to line up, medium rate

Fill cell & Flush 5 times

FXN: READ

COND

Reduce flow to minimum (off?)

DO NOT flush cell

DO NOT move FXN switch away from READ

### FXN: READ

STD

ENTER when STD STANDARDIZE appears

COND NO 0.999 82 from standard bottle

BATCH NO P147 from standard bottle

“ENTER WHEN READY” will appear

Increase flow

When cell is full & sample flowing press ENTER

“Measuring” will appear

“STANDARD 4.22000” will appear, when stable, press COND

COND ends standardization process

SAL take reading

samples

Obtain good cond measurement 3 times:

Rock sample bottle gently

Wipe tube, Install standard sample bottle

Flow to up, medium rate

Fill cell & Flush 3-5 times

FXN: READ wait for ~5 sec to stabilize

COND

Log cond to 5 dec places

Log sal to 4 dec places

FXN: STDBY