SBE Processing configurations used for EN476 Dennis McGillicuddy, Anderson:

CTD001 through CTD11x have Pressure, dual temperature and conductivity sensors and pumps, WET Labs C-Star 25cm Transmissometer, Benthos Altimeter, SBE43 Oxygen 1, Oxygen 2,

WET Labs ECO Fluorometer, Biospherical PAR and Surface PAR sensors and 12 Niskin bottles using the smaller of GSO’s SBE frames.

Files in CTD\ProControl are the ones used during the cruise. 

Files in CTD\ProControl\New\ are the ones which were written on 4Jun10 and used during the cruise. 

476Datcnv_a

Instrument configuration file
D:\CTDData\ctd01.con
X Match instrument configuration file to input file

Input directory
D:\CTDData
Output directory
D:\CTDData\list

X Process scans to end of file. Scans to skip over 0. Binary output, Up& downcast, Create both data and bottle file. Scans marked with bottle confirm bit.. Scan range offset (s) 0. Scan range duration (s) 2.

X Merge separate header file.

Scan Count, Time Elapsed (s), Pressure Digiquartz (db), Depth (saltwater, m) Latitude 18.50 used for all casts (should be reprocessed for 43.50degN to be like EN437 or with exact latitude for each cast to be accurate), Temp 1 (ITS-90, deg C), Temp 2 (ITS-90, deg C), Temp Diff (ITS-90, deg C), Cond 1 (S/m), Cond 2 (S/m), Cond Diff (S/m), V0, Beam Attenuation Wetlab, Beam Transmission Wetlab, V2, Altimeter (m), V3, PAR, V4, Oxygen V SBE43 1, V5, Oxygen V SBE43 2, V7, Fluorescence WETLab ECOFL (mg/m^3), SPAR, Pump Status, Latitude (deg), Longitude (deg).

476Filter_a

“FILTER can be run before CELLTM to remove any residual response time between the temperature and conductivity sensors and to minimize digitization noise.  On average, a low pass filter with a time constant of 0.03 sec reduces the noise in computed salinity slightly.”

for SBE9/11 SBE9plus, older manualSBE4249.txt suggests

Low pass filter B = 0.15s for pressure


Low pass filter A = 0.03s for conductivity

  New SBEDataProcessing_5.32a.pdf suggests - for conductivity and 0.15s for Pressure

Used: B=0.15 for Press & Depth.

Not Used as it was in previous years: A=0.03 for Cond 1, 2 & Diff.

476AlignCTD_a

  Cond

0.073s (secondary only)

  Temp
0s

  DO

4sec (2 at 25C, 5 at 0C),  V4, Oxy V pri, V5, OxyV

476CellTM_a for 9+ w/TC duct & pump running at 3000rpm.

Correct both pri & sec conductivity values

alpha = 0.03

1/beta = 7
LoopEdit (not done for EN476)

  Fixed minimum velocity of 0.25m/s. X Exclude scans marked bad.

SBE Processing configurations used for EN476 Dennis McGillicuddy, Anderson, continued

476Derive_a

  Depth (sw, m Latitude 18.50 Deg N  used for all casts (should be reprocessed for 43.50degN to be like EN437 or with exact latitude for each cast to be accurate. That would be an approx average of the starting Latitudes of all casts. It should be checked and entered by scientific personnel), Salinity 1 & 2 (PSU), Density 1 & 2 (sigma-t, Kg/m^3), Oxygen 1 & 2 (ml/l), Pot Temp 1 & 2 (ITS-90, deg C), Sound Velocity 1 & 2 (Chen-Millero, m/s), Descent Rate (m/s)

476BinAvg_a

  Pressure, 0.5, X Incl #scans/bin, X exclude scans marked bad, skip over 0, Process Downcast only, X Incl surface bin 0, 0, 0
427BottleSum_a

  X Output min/max values for averaged variables, All Averaged Variables, Derived Variables: Salinity 1 & 2 (PSU), Density 1 & 2 (sigma-t, Kg/m^3), Oxygen Saturation (ml/l), Oxygen 1 & 2 (ml/l), Pot Temp 1 & 2 (ITS-90, deg C), Sound Velocity 1 & 2 (Chen-Millero, m/s).

427ASCII_Out_a

  Output header & data files, 60 lines/pg, Label columns top of file, Semicolon, Julian days, All Output variables. Not done: Exclude scans marked bad, Convert elapsed and system time to mm/dd/year hh:mm:ssm Add first column, Replace bad flag.

CTDprocessing476.doc

Page 1 of 2

