## Slide 1: Data Summary for the Moving Vessel Profiler & SeaHorse
Blair Greenan
Bedford Institute of Oceanography
PRISM-RS PIs Meeting
Old Dominion University
11-12 June 2013

## Slide 2
Data Summary
SeaHorse
3 deployments (Eddy3, Ross Bank, Ice Edge Eddy):
  CTD
  Chlorophyll fluorescence
  Downwelling irradiance, Ed, at 4 λs
  O2 concentration & saturation
  ADCP currents
Moving Vessel Profiler (MVP)
  605 profiles
  CTD
  Chlorophyll fluorescence
  Laser optical plankton counter (LOPC)

## Slide 3
MVP CTD downcasts exhibit salinity & density ‘spiking’
Downcast MVP m1 file
Temperature (°C)
Conductivity (S m-1)
Salinity
spike due to differing sensor response times
Upcast MVP m1 file
Temperature (°C)
Conductivity (S m-1)
Salinity
spike absent in corresponding upcast

## Slide 4
Lagging Exercise
Raw data interpolated to a higher frequency: 0.04 m (75 Hz)
Raw conductivity lagged with respect to pressure and temperature: 40 lags tried
Salinity calculated for each lag
Spiking improved, but not completely eliminated
When conductivity leads temperature, negative spikes result (SeaBird Processing notes)
Salinity
Lag Matrix: Lags = 0.04-1.6 m
Pressure (dB)
Lags 21-23 (= 0.84-0.92 m) give best result

## Slide 5
MVP CTD Processing
Examples of downcast salinity and density corrected with lagging and low pass filter

## Slide 6
Moving Vessel Profiler (MVP)
MVP CTD processing completed
Procedure developed to minimize salinity ‘spiking’
Data organized into Matlab structure arrays – one per cast
Posted on PRISM website
LOPC not processed
Summary of BIO Data Status

## Slide 7: Joides Trough to Western Ross Sea MVP Survey
Salinity
Temperature
Fluorescence
LOPC

## Slide 8
SeaHorse
ADCP
ADCP data processed, although some work required to merge upward and downward looking datasets
CTD
CTD data (incl. fluorescence and O2 sat & %) processed and ready to post, subject to very minor QC/QA
We have not yet attempted to compare the O2 sensor against bottle samples
What file format would be most useful for PRISM – NetCDF, Matlab?
Summary of BIO Data Status

## Slide 9
SeaHorse Deployment 3 - Ross Bank
Salinity
Conductivity (S m-1)
Temperature (°C)
Density (kg m-3)
Pressure (dB)
Pressure (dB)
Pressure (dB)
Pressure (dB)

## Slide 10: Ross Bank tidal velocity amplitude (OSU/ESR model)
SeaHorse

## Slide 11
ADCP Deployment 3 - Ross Bank

## Slide 12
SeaHorse cont’d
OCR
OCR data processed
Index of phytoplankton biomass derived from downwelling irradiance sensor (Ed)
Could be used to ‘fill in’ the gaps caused by fluorescence sensor saturation during deployments 2 and 4
Diffuse attenuation coefficient can be easily calculated – required for defining euphotic depth
Summary of BIO Data Status

## Slide 13
Qualitative index of phytoplankton biomass derived from Ed(λ,z)
Example of fluorometer saturation from deployment 2
SeaHorse Chl fluorometer saturated
Chl estimated from Ed(λ,z) measurements (Nahorniak et al., JAOT, 2001. )
Estimates are qualitative at this stage, but could perhaps be refined
Ed derived Chl (approx. μg L-1)
Fluorescence derived Chl (approx. μg L-1)
Depth (m)
Notes: Add some explanation of model.

## Slide 14
There is potential for a study to compare in situ measurements of fluorescence quantum yield with those derived from satellite
Requirements
Extraction of various parameters from satellite data – Dennis and colleagues have provided Matlab scripts for extracting parameters of interest from MODIS Aqua level 3 products
Quite a substantial investment in time required to generate satellite φf
Is there presently sufficient interest/momentum/researcher time?
Satellite Fluorescence Quantum Yield, φf