## Sheet1
PRISM Calibration #2
Calculation of chlorophyll a concentration and standard dilutions for fluorometer calibration
Date(s): 01/05/2011
Calibration performed by: Anna Mosby
Chlorophyll a product #: Sigma C573-1mg
Chlorophyll a lot #: 020M5155V
Calculation of chl a concentration, stock solution:
[chl a], ug/L | = | 868.7883
Abs (A) | = | 0.076167 | (at 664nm)
Extinction coefficient | = | 87.67 | (in 90% acetone)
Cuvette size | = | 1 | (in cm)
Absorbance readings
1 | 0.0762
2 | 0.0765
3 | 0.0758
Expected range of chlorophyll concentrations: 0.01 - 10 ug/L
Calculation of standard dilutions for fluorometer calibration (C1V1 = C2V2)
Bottle # | Initial sample volume (mL) | Initial concentration (ug/L) |  | Final sample volume (mL) | Final concentration (ug/L) | F0 | FA | F0 - FA | Range | K | Average K by Scale
1 | 0.788 | 868.7883 | = | 50 | 13.6921 | 0.885 | 0.384 | 0.501 | high | 27.3295
1 | 0.788 | 868.7883 | = | 50 | 13.6921 | 1.03 | 0.485 | 0.545 |  | 25.1231
1 | 0.788 | 868.7883 | = | 50 | 13.6921 | 0.998 | 0.478 | 0.52 |  | 26.331
2 | 0.33 | 868.7883 | = | 50 | 5.734 | 0.438 | 0.206 | 0.232 | high | 24.7155
2 | 0.33 | 868.7883 | = | 50 | 5.734 | 0.432 | 0.207 | 0.225 |  | 25.4845
2 | 0.33 | 868.7883 | = | 50 | 5.734 | 0.429 | 0.207 | 0.222 |  | 25.8288
3 | 0.06 | 868.7883 | = | 50 | 1.0425 | 0.209 | 0.098 | 0.111 | med | 9.3923
3 | 0.06 | 868.7883 | = | 50 | 1.0425 | 0.205 | 0.099 | 0.106 |  | 9.8353
3 | 0.06 | 868.7883 | = | 50 | 1.0425 | 0.198 | 0.097 | 0.101 |  | 10.3222
4 | 0.03 | 828.105 | = | 50 | 0.49686 | 0.095 | 0.044 | 0.051 | med | 9.7424
4 | 0.03 | 828.105 | = | 50 | 0.49686 | 0.102 | 0.052 | 0.05 |  | 9.9373
4 | 0.03 | 828.105 | = | 50 | 0.49686 | 0.097 | 0.049 | 0.048 |  | 10.3513
Use Bottle #1 to make the dilutions below:
5 | 0.4 | 13.692 | = | 50 | 0.10954 | 0.011 | 0.005 | 0.006 | low | 18.256
5 | 0.4 | 13.692 | = | 50 | 0.10954 | 0.016 | 0.01 | 0.006 |  | 18.256
5 | 0.4 | 13.692 | = | 50 | 0.10954 | 0.006 | 0.004 | 0.002 |  | 54.768
6 | 0.18 | 13.692 | = | 50 | 0.049291 | 0.004 | 0.001 | 0.003 | low | 16.4304
6 | 0.18 | 13.692 | = | 50 | 0.049291 | 0.004 | 0.001 | 0.003 |  | 16.4304
6 | 0.18 | 13.692 | = | 50 | 0.049291 | 0.005 | 0.001 | 0.004 |  | 12.3228
Notes: This machine (I guess also UPS power) is pure evil. However, it does now go down to 0.049. The FS is set at 5 microl as 50% of 9.000 at Medium.

## Sheet2
Calculation of standard dilutions for fluorometer calibration (C1V1 = C2V2)
Bottle # | Initial sample volume (mL) | Initial concentration (ug/L) |  | Final sample volume (mL) | Final concentration (ug/L) | Range
1 | 5 | 1 | = | 50 | 0.1 | High
2 | 3.5 | 1 | = | 50 | 0.07 | High
3 | 0.5 | 1 | = | 50 | 0.01 | High
4 | 0.35 | 1 | = | 50 | 0.007 | Med
5 | 0.25 | 1 | = | 50 | 0.005 | Med
6 | 0.1 | 1 | = | 50 | 0.002 | Med
Use Bottle #2 to make the dilutions below:
7 | 1.25 | 0.07 | = | 50 | 0.00175 | Low
8 | 0.75 | 0.07 | = | 50 | 0.00105 | Low
9 | 0.5 | 0.07 | = | 50 | 0.0007 | Low