## Iodate dispenser
Calibration of Wheaton 10 ml bottle top dispenser used for dispensing standard solution.
Calibration preformed on Dec. 19, 2000 at CSIRO in Hobart
Dispenser was filled with DIW and 10 ml aliquots were dispensed into a small plastic beaker and weighed.
|  |  | rw |  | fbuoy
Temperature of DIW |  | 24 | 0.9973 |  | 1.001
Wt, g
9.9679
9.9638
9.9656
9.9682
9.9708
9.9687
9.9684
9.9684
9.9648
9.9685
|  |  | Vol 24C | Vol 20C
Avg | 9.9675 |  | 10.005 | 10.0046
SD | 0.0020047
|  |  | rw |  | fbuoy
Temperature of DIW |  | 22.9 | 0.99756 |  | 1.001
2003-10-03
SN# 1077 | 9.9852
9.9926
9.9969
9.9888
9.9992
9.9947
9.9944
9.9852
9.9926
9.9865
|  |  | Vol 22.9C | Vol 20C
Avg | 9.9916 |  | 10.0265 | 10.0262
SD | 0.0047018
2003-10-03
SN# 1079 | 9.9725
9.9774
9.9805
9.9809
9.9833
9.9817
9.9833
9.9865
9.9768
9.9813
|  |  | Vol 22.9C | Vol 20C
Avg | 9.9804 |  | 10.0153 | 10.015
SD | 0.0037571
2003-10-03
SN# 1134 | 10.0181
10.0213
10.0205
10.0153
10.02
10.0185
10.0242
10.0146
10.0186
|  |  | Vol 22.9C | Vol 20C
Avg | 10.019 |  | 10.054 | 10.0537
SD | 0.0027859

## Calculations
Oxygen concentration calculations based on WHP Operations and Methods publication of 7/91
written by C. H. Culberson
O2 in ml/L=(((Vx-Vb)*VIO3*NIO3*5598)/(Vstd-Vb))-1000*DOreg)/(Vbot-Vreg)
O2 in umol/L=(((Vx-Vb)*VIO3*NIO3*25000)/(Vstd-Vb))-1000*DOreg)/(Vbot-Vreg)
Vx=endpoint of sample titration in uliters
Vb=reagent blank in uliters
VIO3=volume of IO3 standard in ml
NIO3=normality of iodate standard in equiv/L
Vstd=endpoint of the standard in uliters
DOreg=0.0017 ml O2 or 0.076 umoles O2 in volume of reagents added to samples
Vbot=volume of bottle in ml
Vreg=volume of reagents added (2 ml)

## O2
CTD/Rosette# | Bottle # | Run Date | Sample Date | EP, uL | Vol KIO3, ml | N KIO3 | Std, uL | Blk, uL | Vol reg | BotVol, ml | O2, uM | O2, ml/L
7/21 | 225 |  |  | 1305.73 | 10.026 | 0.01 | 643.6 | -0.1 | 2 | 146.3676 | 351.6831 | 7.8747
7/19 | 226 |  |  | 1300.94 | 10.026 | 0.01 | 643.6 | -0.1 | 2 | 145.5971 | 352.2714 | 7.8878
7/14 | 227 |  |  | 1047.36 | 10.026 | 0.01 | 643.6 | -0.1 | 2 | 145.2589 | 284.1777 | 6.3631
7/12 | 228 |  |  | 839.08 | 10.026 | 0.01 | 643.6 | -0.1 | 2 | 135.9563 | 243.3689 | 5.4494
7/1 | 229 |  |  | 980.87 | 10.026 | 0.01 | 643.6 | -0.1 | 2 | 137.3569 | 281.6402 | 6.3063
| 230 | [8 empty] | 142.9182
| 231 | [8 empty] | 138.8338
| 232 | [8 empty] | 137.5455
| 233 | [8 empty] | 139.7568
| 234 | [8 empty] | 140.5043
| 235 | [8 empty] | 137.7883
| 236 | [8 empty] | 141.4444
| 237 | [8 empty] | 139.6896
| 238 | [8 empty] | 139.2832
| 239 | 2003-11-18 | [7 empty] | 137.6027
| 240 | [8 empty] | 142.239
| 241 | [8 empty] | 142.7798
| 242 | [8 empty] | 142.6594
| 243 | 2003-11-18 | [7 empty] | 140.5053
| 244 | 2003-11-18 | [7 empty] | 142.2731
| 245 | 2003-11-18 | [7 empty] | 139.9775

## Old Palmer bottle volumes
|  |  |  | °C | rw |  |  |  | fbuoy
|  |  | DIW | 24 | 0.9973 |  | Buoyancy correction |  | 1.001
Bottle No. | Empty | Full | Vol (24C) | Vol(20C)
100P | 84.452 | 227.2445 | 143.3295 | 143.3238
101P | 84.952 | 225.3775 | 140.9536 | 140.948
102P | 85.402 | 224.5637 | 139.6851 | 139.6795
103P | 84.2388 | 226.2134 | 142.5085 | 142.5028
104P | 84.2286 | 228.7372 | 145.0521 | 145.0463
105P | 84.1618 | 223.4192 | 139.7811 | 139.7755
106P | 85.9912 | 223.0814 | 137.6058 | 137.6003
107P | 85.0865 | 225.4279 | 140.8692 | 140.8636
108P | 84.7133 | 223.9093 | 139.7195 | 139.7139
109P | 85.5466 | 225.5429 | 140.5228 | 140.5172
110P | 85.4673 | 224.9417 | 139.9989 | 139.9933
111P | 84.5597 | 226.9772 | 142.9531 | 142.9474
112P | 85.2297 | 225.4341 | 140.7317 | 140.726
113P | 83.8293 | 217.5013 | 134.1747 | 134.1693
114P | 83.8244 | 223.8 | 140.502 | 140.4964
115P | 84.54 | 223.0199 | 139.0007 | 138.9951
116P | 86.0119 | 226.9849 | 141.5032 | 141.4975
117P | 85.3855 | 225.3237 | 140.4645 | 140.4589
118P | 85.4802 | 227.9381 | 142.9936 | 142.9879
119P | 85.6576 | 225.9532 | 140.8232 | 140.8176
120P | 83.2334 | 228.08 | 145.3913 | 145.3855
121P | 85.7383 | 226.5056 | 141.2967 | 141.291
122P | 84.5367 | 227.2735 | 143.2736 | 143.2679
123P | 85.4492 | 232.2641 | 147.367 | 147.3611
124P | 84.936 | 234.6342 | 150.2612 | 150.2552
125P | 85.4683 | 227.1948 | 142.2595 | 142.2538
126P | 85.4495 | 223.4206 | 138.49 | 138.4844
127P | 84.6305 | 216.8697 | 132.7365 | 132.7312
128P | 85.1837 | 233.1586 | 148.5314 | 148.5255
129P | 85.6474 | 227.1187 | 142.0033 | 141.9977
130P | 84.5383 | 221.6705 | 137.6479 | 137.6424
131P | 85.008 | 228.2599 | 143.7906 | 143.7849
132P | 86.3757 | 228.0737 | 142.2309 | 142.2252
133P | 85.1368 | 226.8489 | 142.245 | 142.2394
134P | 84.6774 | 223.5234 | 139.3682 | 139.3626
135P | 84.3653 | 220.5615 | 136.7084 | 136.7029
136P | 86.5055 | 225.7143 | 139.7323 | 139.7267
137P | 86.2536 | 228.134 | 142.414 | 142.4083
138P | 84.4294 | 223.1251 | 139.2173 | 139.2117
139P | 87.2229 | 232.7799 | 146.1044 | 146.0986
140P | 85.1274 | 223.0346 | 138.4258 | 138.4203
141P | 88.5494 | 230.9251 | 142.9111 | 142.9054
142P | 86.627 | 226.6021 | 140.5015 | 140.4959
143P | 86.2936 | 238.6043 | 152.8835 | 152.8774
144P | 85.5717 | 237.2192 | 152.2178 | 152.2117
145P | 85.8435 | 228.293 | 142.9852 | 142.9795
146P | 86.0746 | 226.7127 | 141.167 | 141.1614
147P | 86.4491 | 229.2448 | 143.3327 | 143.327
148P | 85.9547 | 224.274 | 138.8395 | 138.8339
149P | 83.113 | 223.9686 | 141.3853 | 141.3797
150P | 84.3232 | 225.3355 | 141.5426 | 141.5369
151P | 86.0976 | 232.9885 | 147.4433 | 147.4374
152P | 84.8629 | 221.8124 | 137.4645 | 137.459
153P | 85.3353 | 225.9817 | 141.1753 | 141.1697
154P | 84.4189 | 226.8484 | 142.9651 | 142.9594
155P | 86.599 | 226.104 | 140.0296 | 140.024
156P | 85.6404 | 226.833 | 141.7236 | 141.7179
157P | 83.7576 | 229.3419 | 146.1318 | 146.126
100 |  |  |  | 144.3167
101 |  |  |  | 145.626
102 |  |  |  | 145.1906
103 |  |  |  | 147.1166
104 |  |  |  | 145.596
105 |  |  |  | 147.7002
106 |  |  |  | 146.0355
107 |  |  |  | 147.0465
108 |  |  |  | 146.1015
109 |  |  |  | 147.7813
110 |  |  |  | 145.4649
111 |  |  |  | 147.6622
112 |  |  |  | 145.6821
113 |  |  |  | 147.0275
114 |  |  |  | 146.6611
115 |  |  |  | 145.5019
116 |  |  |  | 146.0184
117 |  |  |  | 145.3818
118 |  |  |  | 145.2016
119 |  |  |  | 144.9013
120 |  |  |  | 146.0255
121 |  |  |  | 146.2957
122 |  |  |  | 145.1816
123 |  |  |  | 130.8555
124 |  |  |  | 144.1275
125 |  |  |  | 146.7552
126 |  |  |  | 146.2387
200 |  |  |  | 143.9513
201 |  |  |  | 145.8933
202 |  |  |  | 145.6731
203 |  |  |  | 145.2707
204 |  |  |  | 145.0224
205 |  |  |  | 143.7751
206 |  |  |  | 143.5529
207 |  |  |  | 144.8993
208 |  |  |  | 145.2456
209 |  |  |  | 144.4718
210 |  |  |  | 144.9203
211 |  |  |  | 145.7962
212 |  |  |  | 144.7181
213 |  |  |  | 145.2657
214 |  |  |  | 146.2036
215 |  |  |  | 146.57
216 |  |  |  | 146.2477
217 |  |  |  | 145.1615
218 |  |  |  | 146.1496
219 |  |  |  | 146.9154
220 |  |  |  | 145.3317
221 |  |  |  | 145.9774
222 |  |  |  | 145.7572
223 |  |  |  | 146.1836
224 |  |  |  | 146.6481
225 |  |  |  | 146.4559
226 |  |  |  | 146.6841
227 |  |  |  | 146.0034
228 |  |  |  | 146.1666
229 |  |  |  | 145.1225
300 |  |  |  | 145.4639
301 |  |  |  | 146.7762
302 |  |  |  | 143.9463
303 |  |  |  | 144.0804
304 |  |  |  | 143.7351
305 |  |  |  | 144.2686
306 |  |  |  | 144.4969
307 |  |  |  | 144.5569
308 |  |  |  | 143.7801
309 |  |  |  | 145.0715
310 |  |  |  | 145.4829
311 |  |  |  | 143.8422
312 |  |  |  | 145.2136
313 |  |  |  | 145.0294
314 |  |  |  | 145.616
315 |  |  |  | 146.6511
316 |  |  |  | 146.7282
317 |  |  |  | 145.7261
318 |  |  |  | 144.7651
319 |  |  |  | 144.7942
320 |  |  |  | 146.4409
321 |  |  |  | 144.661
322 |  |  |  | 146.8263
323 |  |  |  | 146.1556
324 |  |  |  | 146.6451
325 |  |  |  | 144.2546
326 |  |  |  | 146.7772
327 |  |  |  | 144.9093
328 |  |  |  | 145.577
329 |  |  |  | 146.2737
400 |  |  |  | 142.9182
401 |  |  |  | 144.7751
402 |  |  |  | 145.7602
403 |  |  |  | 145.5179
404 |  |  |  | 145.0024
405 |  |  |  | 144.8863
406 |  |  |  | 145.2486
407 |  |  |  | 145.1776
408 |  |  |  | 143.2005
409 |  |  |  | 144.7191
410 |  |  |  | 145.2196
411 |  |  |  | 143.2405
500 |  |  |  | 144.661
501 |  |  |  | 146.8073
502 |  |  |  | 147.0325
503 |  |  |  | 144.9974
504 |  |  |  | 146.0255
505 |  |  |  | 147.536
506 |  |  |  | 146.7352
507 |  |  |  | 145.635
508 |  |  |  | 146.0124
509 |  |  |  | 146.521
510 |  |  |  | 145.8423
511 |  |  |  | 146.6401
512 |  |  |  | 145.8443
513 |  |  |  | 145.4148
514 |  |  |  | 146.7262
515 |  |  |  | 145.0995
516 |  |  |  | 146.9304
517 |  |  |  | 146.4129
518 |  |  |  | 144.5329
519 |  |  |  | 147.0105
520 |  |  |  | 147.3899
521 |  |  |  | 145.9244
522 |  |  |  | 145.7291
523 |  |  |  | 146.9074
524 |  |  |  | 146.9144
525 |  |  |  | 146.7002
526 |  |  |  | 146.8683
527 |  |  |  | 146.1015
528 |  |  |  | 145.606
529 |  |  |  | 144.619
600 |  |  |  | 147.6101
601 |  |  |  | 144.2516
602 |  |  |  | 145.9934
603 |  |  |  | 146.4649
604 |  |  |  | 145.8483
605 |  |  |  | 146.9915
606 |  |  |  | 145.63
607 |  |  |  | 145.7442
608 |  |  |  | 147.1927
609 |  |  |  | 144.659
610 |  |  |  | 146.0675
611 |  |  |  | 146.4829
612 |  |  |  | 145.7932
613 |  |  |  | 144.7962
614 |  |  |  | 144.8813
615 |  |  |  | 146.0495
616 |  |  |  | 146.503
617 |  |  |  | 146.4659
618 |  |  |  | 146.8123
619 |  |  |  | 146.8994
620 |  |  |  | 146.551
621 |  |  |  | 146.9955
622 |  |  |  | 146.8113
623 |  |  |  | 144.6891
624 |  |  |  | 146.8453
625 |  |  |  | 146.7913
626 |  |  |  | 147.4109
627 |  |  |  | 146.7252
628 |  |  |  | 147.1857
629 |  |  |  | 146.9304
700 |  |  |  | 146.4939
701 |  |  |  | 147.0876
702 |  |  |  | 146.279
703 |  |  |  | 146.0615
704 |  |  |  | 147.1947
705 |  |  |  | 146.2887
706 |  |  |  | 146.7102
707 |  |  |  | 146.6321
708 |  |  |  | 145.6941
709 |  |  |  | 145.3067
710 |  |  |  | 144.5149
711 |  |  |  | 144.7761
712 |  |  |  | 145.6661
713 |  |  |  | 145.6901
714 |  |  |  | 146.0425
715 |  |  |  | 145.2767
716 |  |  |  | 145.0935
717 |  |  |  | 144.8362
718 |  |  |  | 146.4839
719 |  |  |  | 146.5961
720 |  |  |  | 146.4839
721 |  |  |  | 145.581
722 |  |  |  | 145.9254
723 |  |  |  | 146.7352
724 |  |  |  | 147.2507
725 |  |  |  | 145.7522
726 |  |  |  | 144.8522
727 |  |  |  | 146.4299
728 |  |  |  | 146.7692
729 |  |  |  | 145.4979
800 |  |  |  | 146.8403
801 |  |  |  | 144.9904
802 |  |  |  | 145.6901
803 |  |  |  | 146.8904
804 |  |  |  | 147.1056
805 |  |  |  | 147.1526
806 |  |  |  | 146.6271
807 |  |  |  | 146.8884
808 |  |  |  | 144.7221
809 |  |  |  | 146.504
810 |  |  |  | 144.3347
811 |  |  |  | 146.6902
812 |  |  |  | 145.4729
813 |  |  |  | 146.8854
814 |  |  |  | 144.7111
815 |  |  |  | 146.7672
816 |  |  |  | 145.9564
817 |  |  |  | 145.1836
818 |  |  |  | 145.9133
819 |  |  |  | 147.4059
820 |  |  |  | 146.3528
821 |  |  |  | 147.1116
822 |  |  |  | 146.503
823 |  |  |  | 145.9153
824 |  |  |  | 145.6541
825 |  |  |  | 145.713
826 |  |  |  | 144.4888
827 |  |  |  | 144.5609
828 |  |  |  | 146.7202
829 |  |  |  | 146.9785
900 |  |  |  | 146.6151
901 |  |  |  | 146.0335
902 |  |  |  | 144.8773
903 |  |  |  | 145.9524
904 |  |  |  | 145.6881
905 |  |  |  | 147.1406
906 |  |  |  | 146.2697
907 |  |  |  | 147.3208
908 |  |  |  | 147.2087
909 |  |  |  | 145.4859
910 |  |  |  | 144.9423
911 |  |  |  | 146.7813
912 |  |  |  | 147.2928
913 |  |  |  | 147.3288
914 |  |  |  | 145.3788
915 |  |  |  | 146.7522
LMG bottles 10/18/03
200 | 118.726 | 264.226 | 145.9879 | 145.9845 |  |  | °C | rw |  | fbuoy
201 | 115.887 | 263.753 | 148.3618 | 148.3584 |  | DIW | 22.3 | 0.9977 | Buoyancy correction | 1.001
202 | 112.197 | 259.131 | 147.4267 | 147.4233
203 | 115.617 | 259.113 | 143.9772 | 143.9739
204 | 118.244 | 263.67 | 145.9136 | 145.9103
205 | 116.09 | 262.724 | 147.1257 | 147.1223
206 | 117.679 | 266.1156 | 148.9343 | 148.9309
207 | 123.94 | 269.36 | 145.9076 | 145.9043
208 | 119.29 | 264.294 | 145.4902 | 145.4869
209 | 126.268 | 268.785 | 142.9949 | 142.9916
210 | 121.415 | 267.906 | 146.9822 | 146.9788
211 | 126.255 | 272.126 | 146.3601 | 146.3568
212 | 122.591 | 271.083 | 148.9899 | 148.9865
213 | 123.367 | 270.488 | 147.6143 | 147.6109
214 | 123.253 | 269.699 | 146.9371 | 146.9337
215 | 127.307 | 271.787 | 144.9645 | 144.9611
216 | 122.41 | 270.694 | 148.7812 | 148.7778
217 | 126.38 | 270.86 | 144.9645 | 144.9611
218 | 126.112 | 271.344 | 145.719 | 145.7156
219 | 127.254 | 272.177 | 145.409 | 145.4056
220 | 126.25 | 268.602 | 142.8293 | 142.826
221 | 117.237 | 265.73 | 148.9909 | 148.9875
222 | 121.938 | 266.983 | 145.5314 | 145.528
223 | 119.117 | 264.52 | 145.8906 | 145.8872
224 | 111.632 | 257.144 | 145.9999 | 145.9966
225 | 120.721 | 266.606 | 146.3708 | 146.3676 |  |  | °C | rw |  | fbuoy
226 | 125.583 | 270.7 | 145.6003 | 145.5971 |  | DIW | 22.2 | 0.99773 | Buoyancy correction | 1.001
227 | 125.001 | 269.781 | 145.2621 | 145.2589
228 | 84.265 | 219.773 | 135.9593 | 135.9563
229 | 85.946 | 222.85 | 137.3599 | 137.3569
230 | 85.999 | 228.446 | 142.9214 | 142.9182
231 | 86.386 | 224.762 | 138.8368 | 138.8338
232 | 83.869 | 220.961 | 137.5485 | 137.5455
233 | 125.087 | 264.383 | 139.7599 | 139.7568
234 | 126.018 | 266.059 | 140.5074 | 140.5043
235 | 127.763 | 265.097 | 137.7913 | 137.7883
236 | 124.753 | 265.731 | 141.4475 | 141.4444
237 | 131.716 | 270.945 | 139.6926 | 139.6896
238 | 128.929 | 267.753 | 139.2863 | 139.2832
239 | 129.213 | 266.362 | 137.6057 | 137.6027
240 | 125.617 | 267.387 | 142.2421 | 142.239
241 | 121.067 | 263.376 | 142.7829 | 142.7798
242 | 127.813 | 270.002 | 142.6625 | 142.6594
243 | 129.943 | 269.985 | 140.5084 | 140.5053
244 | 126.31 | 268.114 | 142.2762 | 142.2731
245 | 126.85 | 266.366 | 139.9806 | 139.9775
246 | 114.027 | 261.445 | 147.9089 | 147.9057
247 | 118.04 | 266.103 | 148.5561 | 148.5528
248 | 114.121 | 262.756 | 149.13 | 149.1267
249 | 115.728 | 262.373 | 147.1333 | 147.1301
250 | 117.807 | 266.277 | 148.9644 | 148.9611
251 | 117.68 | 264.694 | 147.5036 | 147.5003
252 | 123.65 | 263.399 | 140.2144 | 140.2113
253 | 126.232 | 268.125 | 142.3655 | 142.3624
254 | 131.23 | 269.942 | 139.1739 | 139.1709
255 | 126.051 | 263.062 | 137.4673 | 137.4642
256 | 127.122 | 265.663 | 139.0024 | 138.9993
257 | 126.281 | 266.963 | 141.1505 | 141.1474
258 | 124.486 | 267.548 | 143.5384 | 143.5353
259 | 126.004 | 269.815 | 144.2899 | 144.2867
260 | 124.131 | 270.243 | 146.5986 | 146.5953
261 | 122.713 | 270.179 | 147.9953 | 147.9904 |  |  | °C | rw |  | fbuoy
262 | 124.647 | 270.721 | 146.5983 | 146.5935 |  | DIW | 23.3 | 0.99747 | Buoyancy correction | 1.001
263 | 126.471 | 268.172 | 142.2096 | 142.2049
264 | 122.763 | 269.933 | 147.6983 | 147.6934
265 | 124.779 | 269.102 | 144.841 | 144.8363
266 | 128.418 | 271.626 | 143.722 | 143.7173
267 | 122.77 | 262.764 | 140.4965 | 140.4919
268 | 126.62 | 262.325 | 136.1921 | 136.1876
269 | 130.935 | 271.019 | 140.5868 | 140.5822
270 | 127.276 | 267.619 | 140.8467 | 140.8421
271 | 124.741 | 264.343 | 140.1031 | 140.0985
272 | 124.81 | 262.693 | 138.3779 | 138.3734
273 | 123.957 | 265.504 | 142.0551 | 142.0504
274 | 127.494 | 269.521 | 142.5368 | 142.5321
275 | 126.969 | 264.139 | 137.6624 | 137.6578
276 | 129.209 | 267.687 | 138.9751 | 138.9705
277 | 122.842 | 262.086 | 139.7438 | 139.7392
278 | 126.27 | 266.978 | 141.2131 | 141.2084
279 | 128.349 | 266.3 | 138.4462 | 138.4416
280 | 129.627 | 263.717 | 134.5713 | 134.5669
281 | 127.864 | 268.552 | 141.193 | 141.1883
282 | 125.366 | 267.007 | 142.1494 | 142.1447
283 | 112.51 | 261.787 | 149.8128 | 149.8079
284 | 123.08 | 268.285 | 145.7262 | 145.7214
285 | 109.244 | 259.006 | 150.2996 | 150.2946
286 | 116.95 | 264.094 | 147.6722 | 147.6673
287 | 110.252 | 260.955 | 151.2439 | 151.2389

## Data Logsheet
Oxygen Raw Data Sheet
Date | Std 1 | Std 2 | Std 3 | Std 3 |  |  | R1 | R2 | R2-R1
Sample # | Bottle # | Run Date | Sample Date | EP, uL | Vol KIO3, ml | N KIO3 | Std, uL | Blk, uL | Vol reg | Notes
[5 empty] | 10.029 | 0.01 |  |  | 2
[5 empty] | 10.029 | 0.01 |  |  | 2
[5 empty] | 10.029 | 0.01 |  |  | 2
[5 empty] | 10.029 | 0.01 |  |  | 2
[5 empty] | 10.029 | 0.01 |  |  | 2
[5 empty] | 10.029 | 0.01 |  |  | 2
[5 empty] | 10.029 | 0.01 |  |  | 2
[5 empty] | 10.029 | 0.01 |  |  | 2
[5 empty] | 10.029 | 0.01 |  |  | 2
[5 empty] | 10.029 | 0.01 |  |  | 2
[5 empty] | 10.029 | 0.01 |  |  | 2
[5 empty] | 10.029 | 0.01 |  |  | 2
[5 empty] | 10.029 | 0.01 |  |  | 2
[5 empty] | 10.029 | 0.01 |  |  | 2
[5 empty] | 10.029 | 0.01 |  |  | 2

## BotVols
Latest NBP bottle volumes determined Nov. 11, 2001
Bottle No. | Vol(20C)
100P | 143.6431
109P | 140.5116
110P | 140.0725
111P | 142.8661
113P | 134.076
116P | 141.4549
118P | 142.7706
121P | 141.1454
130P | 137.5906
131P | 143.7478
133P | 142.4271
136P | 139.6755
138P | 139.5138
139P | 145.969
143P | 152.8275
146P | 141.0911
147P | 143.0806
149P | 140.7626
152P | 137.3702
154P | 142.9065
104P | 145.0463 | Bottles in red have not been recalibrated
106P | 137.6003
107P | 140.8636
117P | 140.4589
122P | 143.2679
124P | 150.2552
126P | 138.4844
128P | 148.5255
129P | 141.9977
132P | 142.2252
148P | 138.8339
150P | 141.5369
151P | 147.4374
156P | 141.7179
610 | 146.068 | Bottles provided Hugh Ducklow
707 | 146.632
711 | 144.776
728 | 146.769
910 | 144.942
200 | 145.9845 | LMG bottles 10/18/03
201 | 148.3584
202 | 147.4233
203 | 143.9739
204 | 145.9103
205 | 147.1223
206 | 148.9309
207 | 145.9043
208 | 145.4869
209 | 142.9916
210 | 146.9788
211 | 146.3568
212 | 148.9865
213 | 147.6109
214 | 146.9337
215 | 144.9611
216 | 148.7778
217 | 144.9611
218 | 145.7156
219 | 145.4056
220 | 142.826
221 | 148.9875
222 | 145.528
223 | 145.8872
224 | 145.9966
225 | 146.3676
226 | 145.5971
227 | 145.2589
228 | 135.9563
229 | 137.3569
230 | 142.9182
231 | 138.8338
232 | 137.5455
233 | 139.7568
234 | 140.5043
235 | 137.7883
236 | 141.4444
237 | 139.6896
238 | 139.2832
239 | 137.6027
240 | 142.239
241 | 142.7798
242 | 142.6594
243 | 140.5053
244 | 142.2731
245 | 139.9775
246 | 147.9057
247 | 148.5528
248 | 149.1267
249 | 147.1301
250 | 148.9611
251 | 147.5003
252 | 140.2113
253 | 142.3624
254 | 139.1709
255 | 137.4642
256 | 138.9993
257 | 141.1474
258 | 143.5353
259 | 144.2867
260 | 146.5953
261 | 147.9904
262 | 146.5935
263 | 142.2049
264 | 147.6934
265 | 144.8363
266 | 143.7173
267 | 140.4919
268 | 136.1876
269 | 140.5822
270 | 140.8421
271 | 140.0985
272 | 138.3734
273 | 142.0504
274 | 142.5321
275 | 137.6578
276 | 138.9705
277 | 139.7392
278 | 141.2084
279 | 138.4416
280 | 134.5669
281 | 141.1883
282 | 142.1447
283 | 149.8079
284 | 145.7214
285 | 150.2946
286 | 147.6673
287 | 151.2389
288 | 143.1962
289 | 142.5214
290 | 146.8642
291 | 147.6864
292 | 144.0957
293 | 143.4088
294 | 143.0729
295 | 143.9884
296 | 147.2091
297 | 144.5228
298 | 144.1628
299 | 147.2382
300 | 143.6334
301 | 143.9553
302 | 145.6729
303 | 146.0119
304 | 142.8462
305 | 147.3435
306 | 147.6844
307 | 147.2803
308 | 144.1999
309 | 146.7108
310 | 148.981
311 | 147.5852
312 | 143.869
313 | 143.6514
314 | 144.7194
315 | 146.7048
316 | 145.5667
317 | 147.2523

## Blanks
Reagent blanks
Date | R1 | R2 | Rb
2001-01-24 | 65.1 | 65.6 | -0.5
2001-01-24 | 66.6 | 66.1 | 0.5
2001-01-24 | 65.85 | 65.6 | 0.25
2001-01-24 | 61.85 | 61.6 | 0.25
2001-01-24 | 63.6 | 64.85 | -1.25
|  | Avg | -0.15
|  | SD | 0.6442
Avg | 64.6 | 64.75 | -0.15
SD | 1.6956 | 1.6248

## Standards
Iodate standard was prepared at LDEO on Oct. 20, 2000.
WKIO3 | 0.3567 |  | NIO3(tp) | 0.010002
fbuoy | 1.0002
V(tp) | 1
V(20) | 1
tp | 24
Standards based on assumption that Wheaton dispenser delivered 10.005 ml of standard.
Date | Stds | [5 empty] | Avg | SD
2000-11-22 | 650.98 | 649.67 | 649.83 |  |  |  | 650.16 | 0.5835
2000-12-09 | 653.55 | 649.78 |  |  |  |  | 651.665 | 1.885
2000-12-21 | 643.33 | 643.15 | 644.09 |  |  |  | 643.5233 | 0.40738
2000-12-23 | 657.06* | 642.51 | 642 | 642.25 |  |  | 642.2533 | 0.25502
2000-12-24 | 646.41 | 646.61 | 646.4 | 646.54 |  |  | 646.5167 | 0.10693
2000-12-26 | 643.53 | 644.76 | 645.35 | 645.24 |  |  | 645.1167 | 0.31374
2000-12-26 | 642.45 |  | 644.87 | 642.33 | 646.66 |  | 643.6 | 1.7961
2000-12-29 | 643.59 | 650.4* | 645.58 | 646.62 |  |  | 646.1 | 0.73539
2000-12-30 | 653.14* | 651.2* | 644.07 | 645.77 |  |  | 644.92 | 1.2021
2000-12-31 | 650.2 | 648.75 | 645.59 | 647.21 | 647.94 | 649.19 | 647.1833 | 1.5802
2000-12-31 | 643.85 | 643.95 | 644.81 |  |  |  | 644.2033 | 0.52776
2001-01-01 | 649.18* | 644.08 | 644.63 |  |  |  | 644.355 | 0.38891
2001-01-06 | 642.58 | 644.1 | 647.36* | 642.84 |  |  | 643.34 | 1.0748
New bottle
2001-01-09 | 645.08* | 647.08 | 648.91 | 645.75* | 648.43 | 648.12 | 647.995 | 1.294
2001-01-14 | 653.08 | 648.22 | 656.68* | 650.29 | 651.85 | 649.22 | 650.65 | 3.4365
2001-01-21 | 649.33 | 645.33 | 646.83 | 652.8 |  |  | 648.5725 | 2.8282
646.2596
2.7738
2001-03-22 | 641.08 | 640.24 | 640.92 | 639.81 | 641.67 |  | 640.744 | 0.65234
2003-10-01 | 701.12 | 703.33 | 704.4 | 705.16 |  |  | 703.5025 | 1.5214
2003-10-08 | 703.97 | 701.73 | 702.37 | 704.73 |  |  | 703.2 | 1.2025
2003-10-16 | 702.33 | 702.84 | 702.45 |  |  |  | 702.54 | 0.21772
2003-10-17 |  | 711.33 | 707.27 | 709.29 |  |  | 709.2967 | 1.6575
158
0.6371
597.6