## Iodate dispenser
Calibration of Wheaton 10 ml bottle top dispenser used for dispensing standard solution.
Calibration preformed on Dec. 19, 2000 at CSIRO in Hobart
Dispenser was filled with DIW and 10 ml aliquots were dispensed into a small plastic beaker and weighed.
|  |  | rw |  | fbuoy
Temperature of DIW |  | 24 | 0.9973 |  | 1.001
Wt, g
9.9679
9.9638
9.9656
9.9682
9.9708
9.9687
9.9684
9.9684
9.9648
9.9685
|  |  | Vol 24C | Vol 20C
Avg | 9.9675 |  | 10.005 | 10.0046
SD | 0.0020047

## Calculations
Oxygen concentration calculations based on WHP Operations and Methods publication of 7/91
written by C. H. Culberson
O2 in ml/L=(((Vx-Vb)*VIO3*NIO3*5598)/(Vstd-Vb))-1000*DOreg)/(Vbot-Vreg)
O2 in umol/L=(((Vx-Vb)*VIO3*NIO3*25000)/(Vstd-Vb))-1000*DOreg)/(Vbot-Vreg)
Vx=endpoint of sample titration in uliters
Vb=reagent blank in uliters
VIO3=volume of IO3 standard in ml
NIO3=normality of iodate standard in equiv/L
Vstd=endpoint of the standard in uliters
DOreg=0.0017 ml O2 or 0.076 umoles O2 in volume of reagents added to samples
Vbot=volume of bottle in ml
Vreg=volume of reagents added (2 ml)

## O2
CTD # | Rosette# | Bottle # | Run Date | Sample Date | EP, uL | Std, uL | Blk, uL | BotVol, ml | O2, uM | O2, ml/L | CTD0 | CTD01 | Average
| Test | 499 | 2012-01-05 | 2012-01-04 | 1211.79 | 610.68 | -0.1 | 140.369 | 358.1821 | 8.0202
| Test | 497 | 2012-01-05 | 2012-01-04 | 1236.94 | 610.68 | -0.1 | 142.81 | 359.2885 | 8.045
na | Test | 498 | 2012-01-05 | 2012-01-04 | 1222.25 | 610.68 | -0.1 | 140.754 | 360.2759 | 8.0671
na | Test | 506 | 2012-01-05 | 2012-01-04 | 1231.46 | 610.68 | -0.1 | 141.154 | 361.9512 | 8.1046
na | Test | 502 | 2012-01-05 | 2012-01-04 | 1207.22 | 610.68 | -0.1 | 138.976 | 360.4582 | 8.0712
na | Test | 496 | 2012-01-05 | 2012-01-04 | 1208.59 | 610.68 | -0.1 | 139.48 | 359.5449 | 8.0507
na | Test | 500 | 2012-01-05 | 2012-01-04 | 1232.43 | 610.68 | -0.1 | 142.406 | 359.0066 | 8.0387
na | Test | 507 | 2012-01-05 | 2012-01-04 | 1209.22 | 610.68 | -0.1 | 139.528 | 359.607 | 8.0521
na | Test | 505 | 2012-01-05 | 2012-01-04 | 1216.95 | 610.68 | -0.1 | 139.913 | 360.8989 | 8.081
2 | 1 | 496 | 2012-01-09 | 2012-01-06 | 659.57 | 610.63 | 12.6 | 139.48 | 196.3044 | 4.3955 | 4.008 | 4.0687 | 4.0384
2 | 7 | 497 | 2012-01-09 | 2012-01-06 | 653.34 | 610.63 | 12.6 | 142.81 | 189.8112 | 4.2501 | 3.9047 | 3.9613 | 3.933
2 | 10 | 498 | 2012-01-09 | 2012-01-06 | 2.36 | 610.63 | 12.6 | 140.754 | -3.6349 | -0.081391 | 5.9871 | 6.0752 | 6.0312
2 | 22 | 499 | 2012-01-09 | 2012-01-06 | 1202.51 | 610.63 | 12.6 | 140.369 | 359.1852 | 8.0427 | 7.3887 | 7.528 | 7.4584
3 | 1 | 503 | 2012-01-09 | 2012-01-07 | 665.88 | 610.63 | 12.6 | 140.108 | 197.3231 | 4.4183 | 4.0375 | 4.0949 | 4.0662
3 | 7 | 502 | 2012-01-09 | 2012-01-07 | 623.29 | 610.63 | 12.6 | 138.976 | 185.947 | 4.1636 | 3.8767 | 3.9334 | 3.905
3 | 12 | 501 | 2012-01-09 | 2012-01-07 | 994.93 | 610.63 | 12.6 | 141.123 | 294.823 | 6.6015 | 6.1225 | 6.2703 | 6.1964
3 | 21 | 500 | 2012-01-09 | 2012-01-07 | 1180.9 | 610.63 | 12.6 | 142.406 | 347.5358 | 7.7818 | 7.2844 | 7.4195 | 7.352
4 | 1 | 507 | 2012-01-09 | 2012-01-08 | 452.77 | 610.63 | 12.6 | 139.528 | 133.3337 | 2.9855 | 4.0194 | 4.0773 | 4.0484
4 | 3 | 506 | 2012-01-09 | 2012-01-08 | 570.33 | 610.63 | 12.6 | 141.154 | 167.116 | 3.742 | 3.9672 | 4.0366 | 4.0019
4 | 7 | 505 | 2012-01-09 | 2012-01-08 | 944.17 | 610.63 | 12.6 | 139.913 | 282.0131 | 6.3147 | 6.3674 | 6.4817 | 6.4246
4 | 20 | 504 | 2012-01-09 | 2012-01-08 | 1221.44 | 610.63 | 12.6 | 140.032 | 365.7991 | 8.1908 | 7.5428 | 7.687 | 7.6149
5 | 19 | 510 | 2012-01-11 | 2012-01-09 13:21 | 893.16 | 610.79 | 12.6 | 141.896 | 262.6921 | 5.882 | 7.9947 | 8.1276 | 8.0611
5 | 9 | 511 | 2012-01-11 | 2012-01-09 13:21 | 640.26 | 610.79 | 12.6 | 139.472 | 190.3889 | 4.2631 | 6.8574 | 6.9406 | 6.899
5 | 5 | 512 | 2012-01-11 | 2012-01-09 13:21 | 956.08 | 610.79 | 12.6 | 142.23 | 280.831 | 6.2882 | 5.8484 | 5.9338 | 5.8911
5 | 1 | 513 | 2012-01-11 | 2012-01-09 13:21 | 1003.49 | 610.79 | 12.6 | 140.851 | 297.8995 | 6.6704 | 6.1632 | 6.2453 | 6.2042
6 | 21 | 514 | 2012-01-11 | 2012-01-09 18:44 | 1277.19 | 610.79 | 12.6 | 144.714 | 370.0405 | 8.2857 | 7.7048 | 7.8515 | 7.7782
6 | 10 | 515 | 2012-01-11 | 2012-01-09 18:44 | 1013.1 | 610.79 | 12.6 | 139.101 | 304.6334 | 6.8212 | 6.3034 | 6.3859 | 6.3447
6 | 5 | 516 | 2012-01-11 | 2012-01-09 18:44 | 977.7 | 610.79 | 12.6 | 142.655 | 286.4107 | 6.4131 | 5.9442 | 6.0347 | 5.9894
6 | 1 | 517 | 2012-01-11 | 2012-01-09 18:44 | 917.58 | 610.79 | 12.6 | 141.707 | 270.3575 | 6.0537 | 6.1535 | 6.2371 | 6.1953
7 | 22 | 496 | 2012-01-11 | 2012-01-10 | 1320.5 | 610.79 | 12.6 | 139.48 | 397.3029 | 8.8962 | 8.0976 | 8.2326 | 8.1651
7 | 14 | 497 | 2012-01-11 | 2012-01-10 | 1227.89 | 610.79 | 12.6 | 142.81 | 360.4019 | 8.0699 | 7.4569 | 7.575 | 7.5159
7 | 7 | 498 | 2012-01-11 | 2012-01-10 | 992.71 | 610.79 | 12.6 | 140.754 | 294.8587 | 6.6023 | 6.0864 | 6.1747 | 6.1305
7 | 2 | 499 | 2012-01-11 | 2012-01-10 | 1002.11 | 610.79 | 12.6 | 140.369 | 298.5202 | 6.6843 | 6.1443 | 6.2298 | 6.1871
8 | 22 | 496 | 01/14//12 | 1/12/2012 | 1226.17 | 627.19 | 12.6 | 139.48 | 358.7574 | 8.0331 | 8.3161 | 8.697 | 8.5066
8 | 11 | 497 | 01/14//13 | 1/12/2013 | 1054.7 | 627.19 | 12.6 | 142.81 | 300.7055 | 6.7332 | 6.4331 | 6.5384 | 6.4858
8 | 3 | 498 | 01/14//14 | 1/12/2014 | 1015.24 | 627.19 | 12.6 | 140.754 | 293.5853 | 6.5738 | 6.2867 | 6.38 | 6.3333
8 | 1 | 499 | 01/14//15 | 1/12/2015 | 982.58 | 627.19 | 12.6 | 140.369 | 284.7944 | 6.3769 | 6.0836 | 6.1787 | 6.1311
9 | 1 | 500 | 01/14//16 | 1/13/2012 | 991.84 | 627.19 | 12.6 | 142.406 | 283.3472 | 6.3445 | 6.0317 | 6.1274 | 6.0795
9 | 9 | 501 | 01/14//17 | 1/13/2013 | 990.38 | 627.19 | 12.6 | 141.123 | 285.5331 | 6.3935 | 6.1932 | 6.2909 | 6.2421
9 | 16 | 502 | 01/14//18 | 1/13/2014 | 1123.99 | 627.19 | 12.6 | 138.976 | 329.713 | 7.3827 | 7.0743 | 7.2004 | 7.1373
9 | 20 | 503 | 01/14//19 | 1/13/2015 | 1247.94 | 627.19 | 12.6 | 140.108 | 363.5424 | 8.1402 | 8.1031 | 8.2822 | 8.1926
10 | 22 | 504 | 01/14//20 | 1/14/2012 | 1221.98 | 627.19 | 12.6 | 140.032 | 356.0871 | 7.9733 | 7.6421 | 7.8025 | 7.7223
10 | 16 | 505 | 01/14//21 | 1/14/2013 | 1167.89 | 627.19 | 12.6 | 139.913 | 340.4298 | 7.6227 | 7.391 | 7.6114 | 7.5012
10 | 7 | 506 | 01/14//22 | 1/14/2014 | 1018.83 | 627.19 | 12.6 | 141.154 | 293.7915 | 6.5784 | 6.318 | 6.4371 | 6.3776
10 | 1 | 507 | 01/14//23 | 1/14/2015 | 973.31 | 627.19 | 12.6 | 139.528 | 283.7923 | 6.3545 | 6.0838 | 6.1799 | 6.1319
11 | 21 | 508 | 2012-01-17 | 2012-01-15 | 1373.46 | 620.61 | 12.6 | 142.069 | 399.2092 | 8.9389 | 8.5662 | 8.7588 | 8.6625
11 | 16 | 509 | 2012-01-17 | 2012-01-15 | 1246.37 | 620.61 | 12.6 | 142.294 | 361.2963 | 8.0899 | 7.7818 | 7.9995 | 7.8907
11 | 9 | 510 | 2012-01-17 | 2012-01-15 | 1008.76 | 620.61 | 12.6 | 141.896 | 292.44 | 6.5481 | 6.26 | 6.3638 | 6.3119
11 | 1 | 511 | 2012-01-17 | 2012-01-15 | 965.46 | 620.61 | 12.6 | 139.472 | 284.6369 | 6.3734 | 6.0059 | 6.1075 | 6.0567
20 | 1 | 496 | 2012-01-17 | 2012-01-17 | 922.2 | 620.61 | 12.6 | 139.48 | 271.6734 | 6.0831 | 6.0492 | 6.1163 | 6.0828
20 | 3 | 497 | 2012-01-17 | 2012-01-17 | 1017.25 | 620.61 | 12.6 | 142.81 | 293.0226 | 6.5612 | 6.1972 | 6.267 | 6.2321
20 | 8 | 498 | 2012-01-17 | 2012-01-17 | 1026.47 | 620.61 | 12.6 | 140.754 | 300.0985 | 6.7196 | 6.4172 | 6.4938 | 6.4555
20 | 20 | 499 | 2012-01-17 | 2012-01-17 | 1195.65 | 620.61 | 12.6 | 140.369 | 351.2406 | 7.8648 | 7.7116 | 7.8153 | 7.7634
25 | 21 | 503 | 1/20/12 | 2012-01-17 | 1221.49 | 616.5367 | 12.6 | 140.108 | 362.0318 | 8.1064
25 | 1 | 500 | 2012-01-20 | 2012-01-17 | 979.99 | 616.5367 | 12.6 | 142.406 | 284.8589 | 6.3784
25 | 7 | 501 | 2012-01-20 | 2012-01-17 | 960.9 | 616.5367 | 12.6 | 141.123 | 281.802 | 6.3099
25 | 10 | 502 | 2012-01-20 | 2012-01-17 | 943.88 | 616.5367 | 12.6 | 138.976 | 281.0721 | 6.2936
30 | 22 | 506 | 2012-01-20 | 2012-01-18 | nan | 616.5367 | 12.6 | 141.154 | #VALUE! | #VALUE!
30 | 13 | 507 | 2012-01-20 | 2012-01-18 | nan | 616.5367 | 12.6 | 139.528 | #VALUE! | #VALUE!
30 | 6 | 508 | 2012-01-20 | 2012-01-18 | 1022.85 | 616.5367 | 12.6 | 142.069 | 298.2193 | 6.6775
30 | 1 | 509 | 2012-01-20 | 2012-01-18 | 1000.45 | 616.5367 | 12.6 | 142.294 | 291.1273
37 | 22 | 496 | 2012-01-23 | 2012-01-19 | 1265.98 | 629.0167 | 12.6 | 139.48 | 369.4445 | 8.2724
37 | 10 | 497 | 2012-01-23 | 2012-01-19 | 1067.19 | 629.0167 | 12.6 | 142.81 | 303.4126 | 6.7938
37 | 5 | 498 | 2012-01-23 | 2012-01-19 | 1018.38 | 629.0167 | 12.6 | 140.754 | 293.6321 | 6.5748
37 | 1 | 499 | 2012-01-23 | 2012-01-19 | 992.5 | 629.0167 | 12.6 | 140.369 | 286.8584 | 6.4232
44 | 21 | 506 | 2012-01-23 | 2012-01-21 | 1259.82 | 629.0167 | 12.6 | 141.154 | 363.2036 | 8.1326
44 | 14 | 507 | 2012-01-23 | 2012-01-21 | 1225.78 | 629.0167 | 12.6 | 139.528 | 357.4527 | 8.0039
44 | 3 | 508 | 2012-01-23 | 2012-01-21 | 979.81 | 629.0167 | 12.6 | 142.069 | 279.7 | 6.2629
44 | 1 | 509 | 2012-01-23 | 2012-01-21 | 981.31 | 629.0167 | 12.6 | 142.294 | 279.6853 | 6.2625
51 | 1 | 500 | 2012-01-23 | 2012-01-23 | 1012.79 | 629.0167 | 12.6 | 142.406 | 288.5615 | 6.4613
51 | 4 | 501 | 2012-01-23 | 2012-01-23 | 1003.96 | 629.0167 | 12.6 | 141.123 | 288.6468 | 6.4632
51 | 10 | 502 | 2012-01-23 | 2012-01-23 | 1099.69 | 629.0167 | 12.6 | 138.976 | 321.5345 | 7.1996
51 | 18 | 503 | 2012-01-23 | 2012-01-23 | 1264.13 | 629.0167 | 12.6 | 140.108 | 367.221 | 8.2226

## BotVols
|  |  |  | °C | rw |  | fbuoy
|  |  | DIW | 24 | 0.9973 | Buoyancy correction | 1.001
Bottle No. | Empty | Full | Vol (24C) | Vol(20C)
400 | 84.452 | 227.2445 | 143.3295 | 143.553
401 | 84.952 | 225.3775 | 140.9536 | 141.038
402 | 85.402 | 224.5637 | 139.6851 | 140.136
403 | 84.2388 | 226.2134 | 142.5085 | 144.593
404 | 84.2286 | 228.7372 | 145.0521 | 141.932
405 | 84.1618 | 223.4192 | 139.7811 | 144.056
406 | 85.9912 | 223.0814 | 137.6058 | 141.201
407 | 85.0865 | 225.4279 | 140.8692 | 141.417
408 | 84.7133 | 223.9093 | 139.7195 | 141.33
409 | 85.5466 | 225.5429 | 140.5228 | 142.93
410 | 85.4673 | 224.9417 | 139.9989 | 142.403
411 | 84.5597 | 226.9772 | 142.9531 | 143.468
412 | 85.2297 | 225.4341 | 140.7317 | 142.726
413 | 83.8293 | 217.5013 | 134.1747 | 146.983
414 | 83.8244 | 223.8 | 140.502 | 144.261
415 | 84.54 | 223.0199 | 139.0007 | 143.328
416 | 86.0119 | 226.9849 | 141.5032 | 146.385
417 | 85.3855 | 225.3237 | 140.4645 | 143.849
418 | 85.4802 | 227.9381 | 142.9936 | 140.503
419 | 85.6576 | 225.9532 | 140.8232 | 140.83
420 | 83.2334 | 228.08 | 145.3913 | 142.192
421 | 85.7383 | 226.5056 | 141.2967 | 142.832
422 | 84.5367 | 227.2735 | 143.2736 | 142.837
423 | 85.4492 | 232.2641 | 147.367 | 143.938
424 | 84.936 | 234.6342 | 150.2612 | 142.945
425 | 85.4683 | 227.1948 | 142.2595 | 138.79
426 | 85.4495 | 223.4206 | 138.49 | 141.648
427 | 84.6305 | 216.8697 | 132.7365 | 139.852
428 | 85.1837 | 233.1586 | 148.5314 | 139.396
429 | 85.6474 | 227.1187 | 142.0033 | 0
430 | 84.5383 | 221.6705 | 137.6479 | 145.037
431 | 85.008 | 228.2599 | 143.7906 | 141.374
432 | 86.3757 | 228.0737 | 142.2309 | 142.838
433 | 85.1368 | 226.8489 | 142.245 | 141.014
434 | 84.6774 | 223.5234 | 139.3682 | 144.117
435 | 84.3653 | 220.5615 | 136.7084 | 142.082
436 | 86.5055 | 225.7143 | 139.7323 | 143.338
437 | 86.2536 | 228.134 | 142.414 | 141.798
438 | 84.4294 | 223.1251 | 139.2173 | 143.763
439 | 87.2229 | 232.7799 | 146.1044 | 143.557
440 | 85.1274 | 223.0346 | 138.4258 | 142.425
441 | 88.5494 | 230.9251 | 142.9111 | 136.446
442 | 86.627 | 226.6021 | 140.5015 | 141.227
443 | 86.2936 | 238.6043 | 152.8835 | 137.393
444 | 85.5717 | 237.2192 | 152.2178 | 142.215
445 | 85.8435 | 228.293 | 142.9852 | 143.705
446 | 86.0746 | 226.7127 | 141.167 | 143.701
447 | 86.4491 | 229.2448 | 143.3327 | 137.552
448 | 85.9547 | 224.274 | 138.8395 | 141.986
449 | 83.113 | 223.9686 | 141.3853 | 140.373
450 | 84.3232 | 225.3355 | 141.5426 | 139.818
451 | 86.0976 | 232.9885 | 147.4433 | 141.835
452 | 84.8629 | 221.8124 | 137.4645 | 142.187
453 | 85.3353 | 225.9817 | 141.1753 | 142.47
454 | 84.4189 | 226.8484 | 142.9651 | 139.415
455 | 86.599 | 226.104 | 140.0296 | 140.616
456 | 85.6404 | 226.833 | 141.7236 | 142.219
457 | 83.7576 | 229.3419 | 146.1318 | 143.541
458 |  |  |  | 141.618
459 |  |  |  | 141.2
460 |  |  |  | 141.766
461 |  |  |  | 143.71
462 |  |  |  | 141.22
463 |  |  |  | 141.865
464 |  |  |  | 139.975
465 |  |  |  | 146.29
466 |  |  |  | 143.658
467 |  |  |  | 143.458
468 |  |  |  | 141.651
469 |  |  |  | 140.921
470 |  |  |  | 145.796
471 |  |  |  | 145.015
472 |  |  |  | 140.159
473 |  |  |  | 144.518
474 |  |  |  | 139.935
475 |  |  |  | 143.41
476 |  |  |  | 144.234
477 |  |  |  | 142.847
478 |  |  |  | 149.914
479 |  |  |  | 144.698
480 |  |  |  | 143.159
481 |  |  |  | 142.939
482 |  |  |  | 142.461
483 |  |  |  | 141.97
484 |  |  |  | 142.211
485 |  |  |  | 140.099
486 |  |  |  | 142.98
487 |  |  |  | 142.007
488 |  |  |  | 141.038
489 |  |  |  | 139.925
490 |  |  |  | 143.034
491 |  |  |  | 142.791
492 |  |  |  | 142.866
493 |  |  |  | 142.069
494 |  |  |  | 138.624
495 |  |  |  | 140.371
496 |  |  |  | 139.48
497 |  |  |  | 142.81
498 |  |  |  | 140.754
499 |  |  |  | 140.369
500 |  |  |  | 142.406
501 |  |  |  | 141.123
502 |  |  |  | 138.976
503 |  |  |  | 140.108
504 |  |  |  | 140.032
505 |  |  |  | 139.913
506 |  |  |  | 141.154
507 |  |  |  | 139.528
508 |  |  |  | 142.069
509 |  |  |  | 142.294
510 |  |  |  | 141.896
511 |  |  |  | 139.472
512 |  |  |  | 142.23
513 |  |  |  | 140.851
514 |  |  |  | 144.714
515 |  |  |  | 139.101
516 |  |  |  | 142.655
517 |  |  |  | 141.707
518 |  |  |  | 140.392
519 |  |  |  | 142.947
Note on E3: Using this column for bottle volume

## Blanks
Reagent blanks
Date | R1 | R2 | Rb
2012-01-09 | 82.5 | 69.75 | 12.75
2012-01-09 | 81.96 | 64.64 | 17.32
2012-01-09 | 82.25 | 74.5 | 7.75
|  | Avg | 12.6067
|  | SD | 3.9083
Avg | 82.2367 | 69.63 | 12.6067
SD | 0.22066 | 4.0262

## Standards
Iodate standard was prepared at LDEO on Oct. 20, 2000.
WKIO3 | 0.3567 |  | NIO3(tp) | 0.010002
fbuoy | 1.0002
V(tp) | 1
V(20) | 1
tp | 24
Standards based on assumption that Wheaton dispenser delivered 10.005 ml of standard.
Date | Stds | [5 empty] | Avg | SD
2012-01-09 | 610.51 | 614.61 | 606.78 |  |  |  | 610.6333 | 3.1978
2012-01-11 | 618.13 | 605.32 | 608.91 |  |  |  | 610.7867 | 5.3954
2000-12-21 | 643.33 | 643.15 | 644.09 |  |  |  | 643.5233 | 0.40738
2000-12-23 | 657.06* | 642.51 | 642 | 642.25 |  |  | 642.2533 | 0.25502
2000-12-24 | 646.41 | 646.61 | 646.4 | 646.54 |  |  | 646.5167 | 0.10693
2000-12-26 | 643.53 | 644.76 | 645.35 | 645.24 |  |  | 645.1167 | 0.31374
2000-12-26 | 642.45 |  | 644.87 | 642.33 | 646.66 |  | 643.6 | 1.7961
2000-12-29 | 643.59 | 650.4* | 645.58 | 646.62 |  |  | 646.1 | 0.73539
2000-12-30 | 653.14* | 651.2* | 644.07 | 645.77 |  |  | 644.92 | 1.2021
2000-12-31 | 650.2 | 648.75 | 645.59 | 647.21 | 647.94 | 649.19 | 647.1833 | 1.5802
2000-12-31 | 643.85 | 643.95 | 644.81 |  |  |  | 644.2033 | 0.52776
2001-01-01 | 649.18* | 644.08 | 644.63 |  |  |  | 644.355 | 0.38891
2001-01-06 | 642.58 | 644.1 | 647.36* | 642.84 |  |  | 643.34 | 1.0748
New bottle
2001-01-09 | 645.08* | 647.08 | 648.91 | 645.75* | 648.43 | 648.12 | 647.995 | 1.294
2001-01-14 | 653.08 | 648.22 | 656.68* | 650.29 | 651.85 | 649.22 | 650.65 | 3.4365
2001-01-21 | 649.33 | 645.33 | 646.83 | 652.8 |  |  | 648.5725 | 2.8282
641.2343
11.7318
1/14/2012 standards by Marco Pedulli
date | stds
2012-01-14 | 634.56 | 623.97 | 623.04 | 627.19
2012-01-17 | 621.61 | 619.32 | 620.9 | 620.61
2012-01-20 | 618.81 | 616.58 | 614.22 | 616.5367
2012-01-23 | 632.25 | 628.33 | 626.47 | 629.0167

## Sheet8
100 143.633 | 100 | 143.633 | 100 | 143.633
113 134.054 | 113 | 134.054 | 113 | 134.054
115 138.964 | 115 | 138.964 | 115 | 138.964
116 141.442 | 116 | 141.442 | 116 | 141.442
117 140.408 | 117 | 140.408 | 117 | 140.408
120 145.092 | 120 | 145.092 | 120 | 145.092
122 143.333 | 122 | 143.333 | 122 | 143.333
126 138.540 | 126 | 138.540 | 126 | 138.540
128 148.457 | 128 | 148.457 | 128 | 148.457
132 142.165 | 132 | 142.165 | 132 | 142.165
133 142.347 | 133 | 142.347 | 133 | 142.347
142 140.459 | 142 | 140.459 | 142 | 140.459
147 143.086 | 147 | 143.086 | 147 | 143.086
148 138.772 | 148 | 138.772 | 148 | 138.772
157 146.091 | 157 | 146.091 | 157 | 146.091
200 146.015 | 200 | 146.015 | 200 | 146.015
201 145.768 | 201 | 145.768 | 201 | 145.768
202 147.454 | 202 | 147.454 | 202 | 147.454
203 143.965 | 203 | 143.965 | 203 | 143.965
204 145.929 | 204 | 145.929 | 204 | 145.929
205 146.810 | 205 | 146.810 | 205 | 146.810
206 148.914 | 206 | 148.914 | 206 | 148.914
207 145.856 | 207 | 145.856 | 207 | 145.856
208 145.465 | 208 | 145.465 | 208 | 145.465
209 142.996 | 209 | 142.996 | 209 | 142.996
210 141.986 | 210 | 141.986 | 210 | 141.986
211 146.331 | 211 | 146.331 | 211 | 146.331
212 148.977 | 212 | 148.977 | 212 | 148.977
213 147.626 | 213 | 147.626 | 213 | 147.626
214 146.892 | 214 | 146.892 | 214 | 146.892
215 144.939 | 215 | 144.939 | 215 | 144.939
216 148.761 | 216 | 148.761 | 216 | 148.761
217 144.948 | 217 | 144.948 | 217 | 144.948
218 142.744 | 218 | 142.744 | 218 | 142.744
219 145.443 | 219 | 145.443 | 219 | 145.443
220 142.865 | 220 | 142.865 | 220 | 142.865
221 138.396 | 221 | 138.396 | 221 | 138.396
222 143.858 | 222 | 143.858 | 222 | 143.858
223 145.878 | 223 | 145.878 | 223 | 145.878
224 146.058 | 224 | 146.058 | 224 | 146.058
225 151.492 | 225 | 151.492 | 225 | 151.492
226 145.589 | 226 | 145.589 | 226 | 145.589
227 145.231 | 227 | 145.231 | 227 | 145.231
228 144.345 | 228 | 144.345 | 228 | 144.345
229 137.545 | 229 | 137.545 | 229 | 137.545
230 145.829 | 230 | 145.829 | 230 | 145.829
231 143.283 | 231 | 143.283 | 231 | 143.283
232 146.048 | 232 | 146.048 | 232 | 146.048
233 142.976 | 233 | 142.976 | 233 | 142.976
234 145.823 | 234 | 145.823 | 234 | 145.823
235 138.057 | 235 | 138.057 | 235 | 138.057
236 141.645 | 236 | 141.645 | 236 | 141.645
237 139.631 | 237 | 139.631 | 237 | 139.631
238 139.072 | 238 | 139.072 | 238 | 139.072
239 143.382 | 239 | 143.382 | 239 | 143.382
240 142.222 | 240 | 142.222 | 240 | 142.222
241 143.828 | 241 | 143.828 | 241 | 143.828
242 143.846 | 242 | 143.846 | 242 | 143.846
243 143.953 | 243 | 143.953 | 243 | 143.953
244 142.428 | 244 | 142.428 | 244 | 142.428
245 147.806 | 245 | 147.806 | 245 | 147.806
246 147.873 | 246 | 147.873 | 246 | 147.873
247 146.795 | 247 | 146.795 | 247 | 146.795
248 147.678 | 248 | 147.678 | 248 | 147.678
249 145.720 | 249 | 145.720 | 249 | 145.720
250 146.175 | 250 | 146.175 | 250 | 146.175
251 143.228 | 251 | 143.228 | 251 | 143.228
252 148.258 | 252 | 148.258 | 252 | 148.258
253 147.674 | 253 | 147.674 | 253 | 147.674
254 139.141 | 254 | 139.141 | 254 | 139.141
255 147.834 | 255 | 147.834 | 255 | 147.834
256 139.051 | 256 | 139.051 | 256 | 139.051
257 141.137 | 257 | 141.137 | 257 | 141.137
258 143.516 | 258 | 143.516 | 258 | 143.516
259 144.246 | 259 | 144.246 | 259 | 144.246
260 146.647 | 260 | 146.647 | 260 | 146.647
261 147.991 | 261 | 147.991 | 261 | 147.991
262 146.565 | 262 | 146.565 | 262 | 146.565
263 142.163 | 263 | 142.163 | 263 | 142.163
264 147.647 | 264 | 147.647 | 264 | 147.647
265 144.755 | 265 | 144.755 | 265 | 144.755
266 143.731 | 266 | 143.731 | 266 | 143.731
267 140.478 | 267 | 140.478 | 267 | 140.478
268 136.213 | 268 | 136.213 | 268 | 136.213
269 140.506 | 269 | 140.506 | 269 | 140.506
270 140.852 | 270 | 140.852 | 270 | 140.852
271 140.184 | 271 | 140.184 | 271 | 140.184
272 138.380 | 272 | 138.380 | 272 | 138.380
273 142.966 | 273 | 142.966 | 273 | 142.966
274 142.539 | 274 | 142.539 | 274 | 142.539
275 144.152 | 275 | 144.152 | 275 | 144.152
276 146.732 | 276 | 146.732 | 276 | 146.732
277 139.738 | 277 | 139.738 | 277 | 139.738
278 147.418 | 278 | 147.418 | 278 | 147.418
279 144.768 | 279 | 144.768 | 279 | 144.768
280 134.576 | 280 | 134.576 | 280 | 134.576
281 141.138 | 281 | 141.138 | 281 | 141.138
282 142.166 | 282 | 142.166 | 282 | 142.166
283 142.280 | 283 | 142.280 | 283 | 142.280
284 145.737 | 284 | 145.737 | 284 | 145.737
285 150.782 | 285 | 150.782 | 285 | 150.782
286 147.682 | 286 | 147.682 | 286 | 147.682
287 151.214 | 287 | 151.214 | 287 | 151.214
288 143.974 | 288 | 143.974 | 288 | 143.974
289 147.388 | 289 | 147.388 | 289 | 147.388
290 146.737 | 290 | 146.737 | 290 | 146.737
291 147.523 | 291 | 147.523 | 291 | 147.523
292 144.096 | 292 | 144.096 | 292 | 144.096
293 143.364 | 293 | 143.364 | 293 | 143.364
294 143.044 | 294 | 143.044 | 294 | 143.044
295 144.873 | 295 | 144.873 | 295 | 144.873
296 139.938 | 296 | 139.938 | 296 | 139.938
297 143.176 | 297 | 143.176 | 297 | 143.176
298 144.168 | 298 | 144.168 | 298 | 144.168
299 144.333 | 299 | 144.333 | 299 | 144.333
300 145.267 | 300 | 145.267 | 300 | 145.267
301 143.919 | 301 | 143.919 | 301 | 143.919
302 145.746 | 302 | 145.746 | 302 | 145.746
303 145.960 | 303 | 145.960 | 303 | 145.960
304 143.145 | 304 | 143.145 | 304 | 143.145
305 147.323 | 305 | 147.323 | 305 | 147.323
306 143.876 | 306 | 143.876 | 306 | 143.876
307 147.191 | 307 | 147.191 | 307 | 147.191
308 141.949 | 308 | 141.949 | 308 | 141.949
309 146.364 | 309 | 146.364 | 309 | 146.364
310 144.914 | 310 | 144.914 | 310 | 144.914
311 147.577 | 311 | 147.577 | 311 | 147.577
312 147.505 | 312 | 147.505 | 312 | 147.505
313 140.591 | 313 | 140.591 | 313 | 140.591
314 144.688 | 314 | 144.688 | 314 | 144.688
315 146.658 | 315 | 146.658 | 315 | 146.658
316 148.990 | 316 | 148.990 | 316 | 148.990
317 145.527 | 317 | 145.527 | 317 | 145.527
318 146.132 | 318 | 146.132 | 318 | 146.132
319 141.490 | 319 | 141.490 | 319 | 141.490
320 144.541 | 320 | 144.541 | 320 | 144.541
321 144.372 | 321 | 144.372 | 321 | 144.372
322 143.489 | 322 | 143.489 | 322 | 143.489
323 145.891 | 323 | 145.891 | 323 | 145.891
324 142.974 | 324 | 142.974 | 324 | 142.974
325 143.678 | 325 | 143.678 | 325 | 143.678
326 148.227 | 326 | 148.227 | 326 | 148.227
327 145.198 | 327 | 145.198 | 327 | 145.198
328 139.579 | 328 | 139.579 | 328 | 139.579
329 148.366 | 329 | 148.366 | 329 | 148.366
330 146.575 | 330 | 146.575 | 330 | 146.575
331 147.809 | 331 | 147.809 | 331 | 147.809
332 146.968 | 332 | 146.968 | 332 | 146.968
333 137.932 | 333 | 137.932 | 333 | 137.932
334 140.048 | 334 | 140.048 | 334 | 140.048
335 146.482 | 335 | 146.482 | 335 | 146.482
336 145.469 | 336 | 145.469 | 336 | 145.469
337 146.168 | 337 | 146.168 | 337 | 146.168
338 143.076 | 338 | 143.076 | 338 | 143.076
339 145.262 | 339 | 145.262 | 339 | 145.262
340 143.706 | 340 | 143.706 | 340 | 143.706
341 144.420 | 341 | 144.420 | 341 | 144.420
342 142.889 | 342 | 142.889 | 342 | 142.889
343 146.874 | 343 | 146.874 | 343 | 146.874
344 142.727 | 344 | 142.727 | 344 | 142.727
345 145.433 | 345 | 145.433 | 345 | 145.433
346 143.136 | 346 | 143.136 | 346 | 143.136
347 145.236 | 347 | 145.236 | 347 | 145.236
348 141.954 | 348 | 141.954 | 348 | 141.954
349 141.711 | 349 | 141.711 | 349 | 141.711
350 147.698 | 350 | 147.698 | 350 | 147.698
351 141.515 | 351 | 141.515 | 351 | 141.515
352 143.706 | 352 | 143.706 | 352 | 143.706
353 141.124 | 353 | 141.124 | 353 | 141.124
354 142.810 | 354 | 142.810 | 354 | 142.810
355 143.317 | 355 | 143.317 | 355 | 143.317
356 145.111 | 356 | 145.111 | 356 | 145.111
357 144.242 | 357 | 144.242 | 357 | 144.242
358 144.867 | 358 | 144.867 | 358 | 144.867
359 144.477 | 359 | 144.477 | 359 | 144.477
360 145.188 | 360 | 145.188 | 360 | 145.188
361 148.197 | 361 | 148.197 | 361 | 148.197
362 144.988 | 362 | 144.988 | 362 | 144.988
363 141.320 | 363 | 141.320 | 363 | 141.320
364 148.369 | 364 | 148.369 | 364 | 148.369
365 146.923 | 365 | 146.923 | 365 | 146.923
400 143.553 | 400 | 143.553 | 400 | 143.553
401 141.038 | 401 | 141.038 | 401 | 141.038
402 140.136 | 402 | 140.136 | 402 | 140.136
403 144.593 | 403 | 144.593 | 403 | 144.593
404 141.932 | 404 | 141.932 | 404 | 141.932
405 144.056 | 405 | 144.056 | 405 | 144.056
406 141.201 | 406 | 141.201 | 406 | 141.201
407 141.417 | 407 | 141.417 | 407 | 141.417
408 141.330 | 408 | 141.330 | 408 | 141.330
409 142.930 | 409 | 142.930 | 409 | 142.930
410 142.403 | 410 | 142.403 | 410 | 142.403
411 143.468 | 411 | 143.468 | 411 | 143.468
412 142.726 | 412 | 142.726 | 412 | 142.726
413 146.983 | 413 | 146.983 | 413 | 146.983
414 144.261 | 414 | 144.261 | 414 | 144.261
415 143.328 | 415 | 143.328 | 415 | 143.328
416 146.385 | 416 | 146.385 | 416 | 146.385
417 143.849 | 417 | 143.849 | 417 | 143.849
418 140.503 | 418 | 140.503 | 418 | 140.503
419 140.830 | 419 | 140.830 | 419 | 140.830
420 142.192 | 420 | 142.192 | 420 | 142.192
421 142.832 | 421 | 142.832 | 421 | 142.832
422 142.837 | 422 | 142.837 | 422 | 142.837
423 143.938 | 423 | 143.938 | 423 | 143.938
424 142.945 | 424 | 142.945 | 424 | 142.945
425 138.790 | 425 | 138.790 | 425 | 138.790
426 141.648 | 426 | 141.648 | 426 | 141.648
427 139.852 | 427 | 139.852 | 427 | 139.852
428 139.396 | 428 | 139.396 | 428 | 139.396
430 145.037 | 430 | 145.037 | 430 | 145.037
431 141.374 | 431 | 141.374 | 431 | 141.374
432 142.838 | 432 | 142.838 | 432 | 142.838
433 141.014 | 433 | 141.014 | 433 | 141.014
434 144.117 | 434 | 144.117 | 434 | 144.117
435 142.082 | 435 | 142.082 | 435 | 142.082
436 143.338 | 436 | 143.338 | 436 | 143.338
437 141.798 | 437 | 141.798 | 437 | 141.798
438 143.763 | 438 | 143.763 | 438 | 143.763
439 143.557 | 439 | 143.557 | 439 | 143.557
440 142.425 | 440 | 142.425 | 440 | 142.425
441 136.446 | 441 | 136.446 | 441 | 136.446
442 141.227 | 442 | 141.227 | 442 | 141.227
443 137.393 | 443 | 137.393 | 443 | 137.393
444 142.215 | 444 | 142.215 | 444 | 142.215
445 143.705 | 445 | 143.705 | 445 | 143.705
446 143.701 | 446 | 143.701 | 446 | 143.701
447 137.552 | 447 | 137.552 | 447 | 137.552
448 141.986 | 448 | 141.986 | 448 | 141.986
449 140.373 | 449 | 140.373 | 449 | 140.373
450 139.818 | 450 | 139.818 | 450 | 139.818
451 141.835 | 451 | 141.835 | 451 | 141.835
452 142.187 | 452 | 142.187 | 452 | 142.187
453 142.470 | 453 | 142.470 | 453 | 142.470
454 139.415 | 454 | 139.415 | 454 | 139.415
455 140.616 | 455 | 140.616 | 455 | 140.616
456 142.219 | 456 | 142.219 | 456 | 142.219
457 143.541 | 457 | 143.541 | 457 | 143.541
458 141.618 | 458 | 141.618 | 458 | 141.618
459 141.200 | 459 | 141.200 | 459 | 141.200
460 141.766 | 460 | 141.766 | 460 | 141.766
461 143.710 | 461 | 143.710 | 461 | 143.710
462 141.220 | 462 | 141.220 | 462 | 141.220
463 141.865 | 463 | 141.865 | 463 | 141.865
464 139.975 | 464 | 139.975 | 464 | 139.975
465 146.290 | 465 | 146.290 | 465 | 146.290
466 143.658 | 466 | 143.658 | 466 | 143.658
467 143.458 | 467 | 143.458 | 467 | 143.458
468 141.651 | 468 | 141.651 | 468 | 141.651
469 140.921 | 469 | 140.921 | 469 | 140.921
470 145.796 | 470 | 145.796 | 470 | 145.796
471 145.015 | 471 | 145.015 | 471 | 145.015
472 140.159 | 472 | 140.159 | 472 | 140.159
473 144.518 | 473 | 144.518 | 473 | 144.518
474 139.935 | 474 | 139.935 | 474 | 139.935
475 143.410 | 475 | 143.410 | 475 | 143.410
476 144.234 | 476 | 144.234 | 476 | 144.234
477 142.847 | 477 | 142.847 | 477 | 142.847
478 149.914 | 478 | 149.914 | 478 | 149.914
479 144.698 | 479 | 144.698 | 479 | 144.698
480 143.159 | 480 | 143.159 | 480 | 143.159
481 142.939 | 481 | 142.939 | 481 | 142.939
482 142.461 | 482 | 142.461 | 482 | 142.461
483 141.970 | 483 | 141.970 | 483 | 141.970
484 142.211 | 484 | 142.211 | 484 | 142.211
485 140.099 | 485 | 140.099 | 485 | 140.099
486 142.980 | 486 | 142.980 | 486 | 142.980
487 142.007 | 487 | 142.007 | 487 | 142.007
488 141.038 | 488 | 141.038 | 488 | 141.038
489 139.925 | 489 | 139.925 | 489 | 139.925
490 143.034 | 490 | 143.034 | 490 | 143.034
491 142.791 | 491 | 142.791 | 491 | 142.791
492 142.866 | 492 | 142.866 | 492 | 142.866
493 142.069 | 493 | 142.069 | 493 | 142.069
494 138.624 | 494 | 138.624 | 494 | 138.624
495 140.371 | 495 | 140.371 | 495 | 140.371
496 139.480 | 496 | 139.480 | 496 | 139.480
497 142.810 | 497 | 142.810 | 497 | 142.810
498 140.754 | 498 | 140.754 | 498 | 140.754
499 140.369 | 499 | 140.369 | 499 | 140.369
500 142.406 | 500 | 142.406 | 500 | 142.406
501 141.123 | 501 | 141.123 | 501 | 141.123
502 138.976 | 502 | 138.976 | 502 | 138.976
503 140.108 | 503 | 140.108 | 503 | 140.108
504 140.032 | 504 | 140.032 | 504 | 140.032
505 139.913 | 505 | 139.913 | 505 | 139.913
506 141.154 | 506 | 141.154 | 506 | 141.154
507 139.528 | 507 | 139.528 | 507 | 139.528
508 142.069 | 508 | 142.069 | 508 | 142.069
509 142.294 | 509 | 142.294 | 509 | 142.294
510 141.896 | 510 | 141.896 | 510 | 141.896
511 139.472 | 511 | 139.472 | 511 | 139.472
512 142.230 | 512 | 142.230 | 512 | 142.230
513 140.851 | 513 | 140.851 | 513 | 140.851
514 144.714 | 514 | 144.714 | 514 | 144.714
515 139.101 | 515 | 139.101 | 515 | 139.101
516 142.655 | 516 | 142.655 | 516 | 142.655
517 141.707 | 517 | 141.707 | 517 | 141.707
518 140.392 | 518 | 140.392 | 518 | 140.392
519 142.947 | 519 | 142.947 | 519 | 142.947

## O2 CTD
1 | Bottle | PrDM | Sbeox0ML/L | Sbeox1ML/L | CTD # | Rosette# | O2, uM | O2, ml/L | depth(m) | Oxy-one | Oxy-two | Avg | depth(m)
1 | 1 | 201.09 | 5.5942 | 5.686 | 2 | 1 | 196.3044 | 4.3955 | 703.999 | 4.008 | 4.0687 | 4.0384 | -703.999
1 | 2 | 201.289 | 5.6119 | 5.7012 | 2 | 7 | 189.8112 | 4.2501 | 283.64 | 3.9047 | 3.9613 | 3.933 | -283.64
1 | 3 | 200.743 | 5.6088 | 5.6992 | 2 | 10 |  |  | 101.35 | 5.9871 | 6.0752 | 6.0312 | -101.35
1 | 4 | 201.141 | 5.6032 | 5.7049 | 2 | 22 | 359.1852 | 8.0427 | 10.225 | 7.3887 | 7.528 | 7.4584 | -10.225
1 | 5 | 201.487 | 5.6049 | 5.6978 | 3 | 1 | 197.3231 | 4.4183 | 708.964 | 4.0375 | 4.0949 | 4.0662 | -708.964
1 | 6 | 200.628 | 5.6051 | 5.6964 | 3 | 7 | 185.947 | 4.1636 | 283.217 | 3.8767 | 3.9334 | 3.905 | -283.217
1 | 7 | 201.351 | 5.6138 | 5.7035 | 3 | 12 | 294.823 | 6.6015 | 80.571 | 6.1225 | 6.2703 | 6.1964 | -80.571
1 | 8 | 200.982 | 5.6155 | 5.7016 | 3 | 21 | 347.5358 | 7.7818 | 10.266 | 7.2844 | 7.4195 | 7.352 | -10.266
1 | 9 | 101.309 | 7.3268 | 7.4365 | 4 | 1 | 133.3337 | 2.9855 | 708.568 | 4.0194 | 4.0773 | 4.0484 | -708.568
1 | 10 | 91.151 | 7.3506 | 7.4779 | 4 | 3 | 167.116 | 3.742 | 555.619 | 3.9672 | 4.0366 | 4.0019 | -555.619
1 | 11 | 81.214 | 7.391 | 7.5078 | 4 | 7 | 282.0131 | 6.3147 | 202.305 | 6.3674 | 6.4817 | 6.4246 | -202.305
1 | 12 | 71.066 | 7.4401 | 7.5445 | 4 | 20 | 365.7991 | 8.1908 | 9.547 | 7.5428 | 7.687 | 7.6149 | -9.547
1 | 13 | 70.691 | 7.4512 | 7.5588 | 5 | 19 | 262.6921 | 5.882 | 10.489 | 7.9947 | 8.1276 | 8.0611 | -10.489
1 | 14 | 60.443 | 7.4825 | 7.5905 | 5 | 9 | 190.3889 | 4.2631 | 81.133 | 6.8574 | 6.9406 | 6.899 | -81.133
1 | 15 | 49.719 | 7.4663 | 7.5939 | 5 | 5 | 280.831 | 6.2882 | 253.042 | 5.8484 | 5.9338 | 5.8911 | -253.042
1 | 16 | 40.343 | 7.4708 | 7.6045 | 5 | 1 | 297.8995 | 6.6704 | 566.119 | 6.1632 | 6.2453 | 6.2042 | -566.119
1 | 17 | 40.438 | 7.4652 | 7.6093 | 6 | 21 | 370.0405 | 8.2857 | 10.007 | 7.7048 | 7.8515 | 7.7782 | -10.007
1 | 18 | 30.461 | 7.4779 | 7.6189 | 6 | 10 | 304.6334 | 6.8212 | 100.807 | 6.3034 | 6.3859 | 6.3447 | -100.807
1 | 19 | 21.247 | 7.4668 | 7.6056 | 6 | 5 | 286.4107 | 6.4131 | 303.031 | 5.9442 | 6.0347 | 5.9894 | -303.031
1 | 20 | 21.242 | 7.467 | 7.6114 | 6 | 1 | 270.3575 | 6.0537 | 504.941 | 6.1535 | 6.2371 | 6.1953 | -504.941
1 | 21 | 10.284 | 7.4449 | 7.5619 | 7 | 22 | 397.3029 | 8.8962 | 10.031 | 8.0976 | 8.2326 | 8.1651 | -10.031
1 | 22 | 10.248 | 7.4288 | 7.5555 | 7 | 14 | 360.4019 | 8.0699 | 49.802 | 7.4569 | 7.575 | 7.5159 | -49.802
1 | 23 | 0.778 | 7.3811 | 7.5109 | 7 | 7 | 294.8587 | 6.6023 | 272.779 | 6.0864 | 6.1747 | 6.1305 | -272.779
1 | 24 | 0.832 | 7.4155 | 7.5337 | 7 | 2 | 298.5202 | 6.6843 | 505.979 | 6.1443 | 6.2298 | 6.1871 | -505.979
2 | 1 | 703.999 | 4.008 | 4.0687 | 8 | 22 | 358.7574 | 8.0331 | 9.739 | 8.3161 | 8.697 | 8.5066 | -9.739
2 | 2 | 604.387 | 4.0054 | 4.071 | 8 | 11 | 300.7055 | 6.7332 | 100.907 | 6.4331 | 6.5384 | 6.4858 | -100.907
2 | 3 | 503.729 | 4.0003 | 4.0616 | 8 | 3 | 293.5853 | 6.5738 | 627.687 | 6.2867 | 6.38 | 6.3333 | -627.687
2 | 4 | 403.566 | 3.9727 | 4.0347 | 8 | 1 | 284.7944 | 6.3769 | 817.553 | 6.0836 | 6.1787 | 6.1311 | -817.553
2 | 5 | 403.433 | 3.9689 | 4.0337 | 9 | 1 | 283.3472 | 6.3445 | 687.194 | 6.0317 | 6.1274 | 6.0795 | -687.194
2 | 6 | 302.516 | 3.9133 | 3.9721 | 9 | 9 | 285.5331 | 6.3935 | 201.67 | 6.1932 | 6.2909 | 6.2421 | -201.67
2 | 7 | 283.64 | 3.9047 | 3.9613 | 9 | 16 | 329.713 | 7.3827 | 39.966 | 7.0743 | 7.2004 | 7.1373 | -39.966
2 | 8 | 202.487 | 4.3016 | 4.3564 | 9 | 20 | 363.5424 | 8.1402 | 9.72 | 8.1031 | 8.2822 | 8.1926 | -9.72
2 | 9 | 152.669 | 5.0652 | 5.1707 | 10 | 22 | 356.0871 | 7.9733 | 9.902 | 7.6421 | 7.8025 | 7.7223 | -9.902
2 | 10 | 101.35 | 5.9871 | 6.0752 | 10 | 16 | 340.4298 | 7.6227 | 40.777 | 7.391 | 7.6114 | 7.5012 | -40.777
2 | 11 | 80.688 | 6.2398 | 6.3303 | 10 | 7 | 293.7915 | 6.5784 | 192.48 | 6.318 | 6.4371 | 6.3776 | -192.48
2 | 12 | 70.596 | 6.2871 | 6.4048 | 10 | 1 | 283.7923 | 6.3545 | 437.63 | 6.0838 | 6.1799 | 6.1319 | -437.63
2 | 13 | 60.281 | 6.4184 | 6.5111
2 | 14 | 60.291 | 6.3622 | 6.4556
2 | 15 | 50.232 | 6.5846 | 6.6874
2 | 16 | 50.098 | 6.6033 | 6.771
2 | 17 | 39.59 | 7.0165 | 7.0772
2 | 18 | 30.743 | 7.3854 | 7.5221
2 | 19 | 20.562 | 7.3904 | 7.5177
2 | 20 | 20.763 | 7.3907 | 7.5193
2 | 21 | 10.32 | 7.3915 | 7.5211
2 | 22 | 10.225 | 7.3887 | 7.528
2 | 23 | 0.816 | 7.4023 | 7.5307
2 | 24 | 1.28 | 7.3971 | 7.521
3 | 1 | 708.964 | 4.0375 | 4.0949
3 | 2 | 607.349 | 4.0299 | 4.0913
3 | 3 | 506.11 | 4.0058 | 4.067
3 | 4 | 404.637 | 3.951 | 4.0108
3 | 5 | 404.597 | 3.949 | 4.0085
3 | 6 | 303.237 | 3.882 | 3.9353
3 | 7 | 283.217 | 3.8767 | 3.9334
3 | 8 | 202.314 | 4.2094 | 4.2859
3 | 9 | 151.47 | 4.7708 | 4.8351
3 | 10 | 101.053 | 5.5778 | 5.6981
3 | 11 | 80.588 | 6.1115 | 6.2459
3 | 12 | 80.571 | 6.1225 | 6.2703
3 | 13 | 70.489 | 6.2457 | 6.3429
3 | 14 | 60.236 | 6.3411 | 6.4359
3 | 15 | 50.765 | 6.3895 | 6.4889
3 | 16 | 40.604 | 6.4609 | 6.5491
3 | 17 | 40.627 | 6.4793 | 6.5631
3 | 18 | 30.264 | 6.6028 | 6.7068
3 | 19 | 20.24 | 7.0648 | 7.315
3 | 20 | 20.284 | 7.1114 | 7.3205
3 | 21 | 10.266 | 7.2844 | 7.4195
3 | 22 | 10.26 | 7.2997 | 7.4161
3 | 23 | 0.952 | 7.3076 | 7.4225
3 | 24 | 0.96 | 7.3065 | 7.423
4 | 1 | 708.568 | 4.0194 | 4.0773
4 | 2 | 606.876 | 3.9805 | 4.0555
4 | 3 | 555.619 | 3.9672 | 4.0366
4 | 4 | 505.759 | 3.9848 | 4.0489
4 | 5 | 404.667 | 4.3108 | 4.3878
4 | 6 | 303.075 | 4.9502 | 5.0215
4 | 7 | 202.305 | 6.3674 | 6.4817
4 | 8 | 151.116 | 6.4 | 6.5054
4 | 9 | 99.924 | 6.3977 | 6.5045
4 | 10 | 80.857 | 6.4504 | 6.5626
4 | 11 | 70.466 | 6.4561 | 6.5632
4 | 12 | 60.096 | 6.4874 | 6.5885
4 | 13 | 50.486 | 6.7153 | 6.8065
4 | 14 | 39.248 | 7.042 | 7.1587
4 | 15 | 39.626 | 7.0506 | 7.1703
4 | 16 | 40.149 | 7.0335 | 7.1764
4 | 17 | 30.21 | 7.3927 | 7.5246
4 | 18 | 19.797 | 7.5421 | 7.6787
4 | 19 | 20.321 | 7.5336 | 7.672
4 | 20 | 9.547 | 7.5428 | 7.687
4 | 21 | 9.373 | 7.5509 | 7.688
4 | 22 | 9.87 | 7.5545 | 7.6865
4 | 23 | 1.285 | 7.5565 | 7.6794
4 | 24 | 1.476 | 7.5515 | 7.6741
5 | 1 | 566.119 | 6.1632 | 6.2453
5 | 2 | 505.896 | 6.1661 | 6.2464
5 | 3 | 404.705 | 6.1144 | 6.201
5 | 4 | 303.075 | 6.0998 | 6.1917
5 | 5 | 253.042 | 5.8484 | 5.9338
5 | 6 | 202.485 | 6.0122 | 6.0938
5 | 7 | 151.583 | 6.0649 | 6.1471
5 | 8 | 101.287 | 6.5438 | 6.6225
5 | 9 | 81.133 | 6.8574 | 6.9406
5 | 10 | 71.037 | 6.7352 | 6.8288
5 | 11 | 60.84 | 7.0423 | 7.2275
5 | 12 | 50.685 | 7.3521 | 7.5412
5 | 13 | 40.995 | 7.911 | 8.0544
5 | 14 | 30.572 | 7.9735 | 8.1079
5 | 15 | 25.437 | 7.9867 | 8.1122
5 | 16 | 25.376 | 7.9901 | 8.1166
5 | 17 | 20.544 | 7.9894 | 8.1301
5 | 18 | 20.736 | 7.9868 | 8.1197
5 | 19 | 10.489 | 7.9947 | 8.1276
5 | 20 | 10.536 | 7.9907 | 8.1387
5 | 21 | 10.478 | 7.9988 | 8.1421
5 | 22 | 10.495 | 8.0068 | 8.1337
5 | 23 | 1.348 | 8.0102 | 8.1458
5 | 24 | 1.404 | 8.0081 | 8.1469
6 | 1 | 504.941 | 6.1535 | 6.2371
6 | 2 | 505.029 | 6.1517 | 6.2369
6 | 3 | 405.191 | 6.0024 | 6.0922
6 | 4 | 405.369 | 6.0048 | 6.0802
6 | 5 | 303.031 | 5.9442 | 6.0347
6 | 6 | 303.2 | 5.9338 | 6.0263
6 | 7 | 302.863 | 5.9448 | 6.0285
6 | 8 | 202.237 | 6.1978 | 6.2872
6 | 9 | 151.258 | 6.2035 | 6.3046
6 | 10 | 100.807 | 6.3034 | 6.3859
6 | 11 | 80.801 | 6.4048 | 6.5037
6 | 12 | 70.345 | 6.3459 | 6.4453
6 | 13 | 60.195 | 6.4046 | 6.5351
6 | 14 | 50.05 | 6.4857 | 6.5867
6 | 15 | 40.634 | 6.6385 | 6.7937
6 | 16 | 40.517 | 6.6891 | 6.8445
6 | 17 | 30.28 | 7.0107 | 7.1633
6 | 18 | 30.333 | 6.899 | 7.226
6 | 19 | 20.033 | 7.4978 | 7.6206
6 | 20 | 20.13 | 7.503 | 7.6871
6 | 21 | 10.007 | 7.7048 | 7.8515
6 | 22 | 9.829 | 7.7007 | 7.8375
6 | 23 | 0.598 | 7.745 | 7.8741
6 | 24 | 0.312 | 7.7512 | 7.869
7 | 1 | 563.693 | 6.1387 | 6.2253
7 | 2 | 505.979 | 6.1443 | 6.2298
7 | 3 | 506.02 | 6.1487 | 6.2353
7 | 4 | 404.719 | 6.1062 | 6.1882
7 | 5 | 404.718 | 6.1072 | 6.1914
7 | 6 | 303.116 | 6.2187 | 6.3054
7 | 7 | 272.779 | 6.0864 | 6.1747
7 | 8 | 201.913 | 6.2284 | 6.3185
7 | 9 | 151.411 | 6.3381 | 6.429
7 | 10 | 100.493 | 6.5323 | 6.6257
7 | 11 | 80.499 | 6.7872 | 6.8896
7 | 12 | 70.774 | 7.1688 | 7.2928
7 | 13 | 59.992 | 7.2925 | 7.4172
7 | 14 | 49.802 | 7.4569 | 7.575
7 | 15 | 40.1 | 7.6006 | 7.7411
7 | 16 | 30.018 | 7.7199 | 7.8432
7 | 17 | 25.065 | 7.8328 | 7.9351
7 | 18 | 25.035 | 7.8252 | 7.9569
7 | 19 | 19.507 | 7.9375 | 8.0426
7 | 20 | 14.696 | 8.0259 | 8.163
7 | 21 | 10.042 | 8.0825 | 8.2259
7 | 22 | 10.031 | 8.0976 | 8.2326
7 | 23 | 1.734 | 8.1155 | 8.243
7 | 24 | 1.685 | 8.1116 | 8.2434
8 | 1 | 817.553 | 6.0836 | 6.1787
8 | 2 | 708.833 | 6.1926 | 6.2822
8 | 3 | 627.687 | 6.2867 | 6.38
8 | 4 | 607.418 | 6.3513 | 6.438
8 | 5 | 505.928 | 6.3616 | 6.4603
8 | 6 | 404.377 | 6.3458 | 6.4049
8 | 7 | 404.397 | 6.3708 | 6.4195
8 | 8 | 303.401 | 6.3124 | 6.4071
8 | 9 | 202.253 | 6.2763 | 6.3701
8 | 10 | 151.363 | 6.2806 | 6.3754
8 | 11 | 100.907 | 6.4331 | 6.5384
8 | 12 | 80.812 | 6.476 | 6.578
8 | 13 | 70.584 | 6.431 | 6.432
8 | 14 | 60.312 | 6.3996 | 6.4606
8 | 15 | 50.251 | 6.3663 | 6.493
8 | 16 | 40.211 | 6.6552 | 6.8263
8 | 17 | 29.835 | 7.0676 | 7.2113
8 | 18 | 19.747 | 7.3186 | 7.4639
8 | 19 | 19.806 | 7.4164 | 7.5999
8 | 20 | 9.719 | 8.1401 | 8.678
8 | 21 | 9.689 | 8.3952 | 8.7212
8 | 22 | 9.739 | 8.3161 | 8.697
8 | 23 | 1.291 | 8.5039 | 8.6508
8 | 24 | 1.342 | 8.5073 | 8.6434
9 | 1 | 687.194 | 6.0317 | 6.1274
9 | 2 | 607.157 | 6.1439 | 6.2659
9 | 3 | 607.194 | 6.1662 | 6.2754
9 | 4 | 505.716 | 6.1897 | 6.2977
9 | 5 | 404.54 | 6.2831 | 6.3952
9 | 6 | 404.537 | 6.2887 | 6.3893
9 | 7 | 302.999 | 6.2717 | 6.3723
9 | 8 | 302.961 | 6.2693 | 6.3694
9 | 9 | 201.67 | 6.1932 | 6.2909
9 | 10 | 151.578 | 6.2954 | 6.4102
9 | 11 | 100.685 | 6.3427 | 6.451
9 | 12 | 80.668 | 6.4175 | 6.537
9 | 13 | 70.628 | 6.4496 | 6.5616
9 | 14 | 60.09 | 6.4644 | 6.5719
9 | 15 | 49.89 | 6.5631 | 6.6845
9 | 16 | 39.966 | 7.0743 | 7.2004
9 | 17 | 29.828 | 7.2387 | 7.3306
9 | 18 | 19.814 | 7.5802 | 7.7759
9 | 19 | 19.826 | 7.5456 | 7.7045
9 | 20 | 9.72 | 8.1031 | 8.2822
9 | 21 | 9.743 | 8.1016 | 8.2904
9 | 22 | 9.725 | 8.0881 | 8.2698
9 | 23 | 1.17 | 8.0988 | 8.2721
9 | 24 | 1.243 | 8.0928 | 8.2708
10 | 1 | 437.63 | 6.0838 | 6.1799
10 | 2 | 404.348 | 6.1141 | 6.2185
10 | 3 | 404.369 | 6.1084 | 6.2192
10 | 4 | 404.36 | 6.1206 | 6.2173
10 | 5 | 303.038 | 6.1273 | 6.2505
10 | 6 | 202.452 | 6.2444 | 6.3765
10 | 7 | 192.48 | 6.318 | 6.4371
10 | 8 | 152.043 | 6.1635 | 6.2953
10 | 9 | 101.316 | 6.1825 | 6.3005
10 | 10 | 81 | 6.1929 | 6.3137
10 | 11 | 70.929 | 6.302 | 6.4222
10 | 12 | 60.891 | 6.5446 | 6.7704
10 | 13 | 50.364 | 6.9928 | 7.2033
10 | 14 | 40.752 | 7.3815 | 7.6288
10 | 15 | 40.762 | 7.3525 | 7.6007
10 | 16 | 40.777 | 7.391 | 7.6114
10 | 17 | 30.508 | 7.6976 | 7.894
10 | 18 | 20.503 | 7.6742 | 7.8423
10 | 19 | 20.508 | 7.6956 | 7.8423
10 | 20 | 9.891 | 7.647 | 7.8063
10 | 21 | 9.897 | 7.6475 | 7.8041
10 | 22 | 9.902 | 7.6421 | 7.8025
10 | 23 | 1.761 | 7.6357 | 7.8011
10 | 24 | 1.777 | 7.6452 | 7.7953