## QSR2200 cal 20313 3-24-08
Calibration Date |  | 2008-03-24
Model Number |  | QSR-2200
Serial Number |  | 20313 | with top on
Operator |  | TPC
Standard Lamp |  | 91537(10/25/2006)
Probe Excitation Voltage Range: |  |  | 6 | to | 18 | VDC(+)
Output Polarity: |  | Positive
Probe Conditions at Calibration(in air):
| Calibration Voltage: |  | 6 |  | VDC(+)
| Probe Current: |  | 3.734 |  | mA
Probe Output Voltage:
| Probe Illuminated |  | 88.09 |  | mV
| Probe Dark |  | 1.925 |  | mV
| Probe Net Response |  | 86.165 |  | mV
| RG780 |  | 2.52 |  | mV
Corrected Lamp Output:
Output In Air (same condition as calibration):
|  | 8834939065935320 | quanta/cm²sec
|  | 0.014671 | uE/cm2sec
Calibration Scale Factor:
(To calculate irradiance, divide the net voltage reading in Volts by this value.)
| Dry: | 9.7528e-18 | V/(quanta/cm²sec)
|  | 5.8731 | V/(uE/cm2sec)
Notes:
1. Annual calibration is recommended.
2. Calibration is performed using a Standard of Spectral Irradiance traceable to the
National Institute of Standards and Technology (NIST).
3. The collector should be cleaned frequently with alcohol.
4. Calibration was performed with customer cable, when available.