CTD SYSTEM CONFIGURATION			Cruise:	EN437	Dates:	21 Jun-05 Jul, 2007

Tech: Lynne Butler					Chief Scientist, Institution: Dennis McGillicuddy, WHOI

Program, versions					SeaSave Win32 v 5.37d, DOS Seasoft 4.249,

# SBE Processing v5.32a

File names						EN437_0607_12pl.cfg  CTD001-CTD108.*

EN437_0607_12pl0343.cfg  CTD109-CTD187.*

Batch files used	pro437.bat & wpro437_a.txt to process, arc437.bat to archive

Deck cal files		EN437\CTD\sswin\CalData\EN437McGillicuddy\  437xapooFs1_0607.dat & 437xapooFs2_0607.dat

SBE 9plus CTD underwater unit (6800m depth capability)	S/N	12293-0444 (07-Mar-03), 23539-0607 (03-Mar-06), none

Pressure sensor, Digiquartz (with temp comp, SBE)		S/N	64853 (07-Mar-03), 80016 (03-Mar-06)

SBE 11plus deck unit					S/N	11P12293- 0426, 0427 (21-Mar-97), none

Temperature sensor, primary (3400 m, 6800m, slow, fast)	S/N	604, 753, 1164, 1185, 1678,

(SBE 3, 3Plus)				2034, 2107, 2899, 2902, 4126, 4130 (13-Dec-06), 4333

Temperature sensor, secondary (3400 m, 6800m, slow, fast)	S/N	604, 753, 1164, 1185, 1678, 2034, 2107, 2899

(SBE 3, 3Plus)				2902, 4126 (13-Dec-06), 4130, 4333, none

Conductivity sensor, primary (SBE 4C, 3400m, 6800m)	S/N	199, 200, 618, 856, 864, 1446 (13-Dec-06),

1578, 1745, 1749, 2459, 2469, 2822

Conductivity sensor, secondary (SBE 4C, 3400m, 6800m)	S/N	199, 200, 618, 856, 864, 1446, 1578, 1745

1749, 2459, 2469, 2822, 3220 (01-Nov-06), none

Pump, primary	(SBE 5T, 10500m, 3000RPM)		S/N	1413, 1595, 1596 (30-Dec-04), 2237, 2238,

2825, 2892 (30-Dec-04)

Pump, secondary	(SBE 5T, 10500m, 3000RPM)		S/N	1413, 1595, 1596, 2237, 2238,

2825 (30-Dec-04), 2892 (30-Dec-04), none

Oxygen sensor, primary (SBE 13, SBE 23, SBE 43, 6800m)	S/N	430343, 430345, 430348, 430372,

DO sensor Serial Number 1230 is brand new				431230 (05-Jun-07p), none

Oxygen sensor, secondary (SBE 13, SBE 23, SBE 43, 6800m)	S/N	430497 (08-Jun-07p, CTD001-108)

430343 (03-Jan-07p CTD109-211)

DO sensor Serial Number 0497 is a loaner from SBE to replace S/N 0348. 0348 went back to SBE for post cal & although it was working very well during EN435, was found to have a perforated membrane & post cal was not performed.

DO sensor Serial Number 0497 had increased values from CTD053 downcast until it was removed from package after CTD108. Probable perforated membrane

DO sensor Serial Number 0343 was last used with the MOCNESS on R/V Kilo Moana.

Fluorometer, primary	(Sea Tech flashlamp, 3000m,	S/N	117S (30Nov04 10microg/l full scale), V6=4.923

WET Labs ECOFluor, 6000m)		FLRTD-231 (18-Dec-06), none

Fluorometer, secondary	(Sea Tech flashlamp, 3000m)	S/N	117S, FLRTD-231, none

WET Labs ECOFluor, 6000m)

Transmissometer (WET Labs C-Star, 25 cm path, 6000m)	S/N	CST-399DR, CST-480DR (back from factory

08-Mar-05 & 21-Dec-01, deck cal 16-May-07&20Jun07),

CST-593DR, none

Altimeter (Benthos PSA-916, 2500m)	deck V2=4.926	S/N	1017 (new Nov-02), 1075 (ship cal 08-Apr-06,

new 22-Oct-03), none

Light sensor PAR, QSP-200L4S/2300 (Biospherical, 1000m)	S/N	4479 (08-Feb-07), 4701(24-Feb-06), none

Light sensor, Surface SPAR, QSR-2200 (Biospherical)	S/N	20121 (24-Feb-06), 20190 (18-Feb-07), none

CTD SYSTEM CONFIGURATION, Continued

Data format (factory & user installed sensors) See also EN437\CTD\CTDReadMe437.txt

Frequency 0		temperature, primary 4130		External voltage 3 1.941    PAR 4479

Frequency 1		conductivity, primary 1446	External voltage 4 1.355	DO 1230

Frequency 2		pressure 80016			External voltage 5 1.402	DO 0497

Frequency 3		temperature, secondary 4126	Ext voltage 6	 4.923	spare

Frequency 4		conductivity, secondary 3220	Ext voltage 7	 1.538	Fluor 231

External voltage 0 4.8083	Transmissometer 480DR		External voltage 8	spare

External voltage 1 0.000	spare				External voltage 9 4.371	SPAR 20190

External voltage 2 4.926	Alt 1075				Pump1=2892, Pump2=2825

Carousel	 pylon, rosette, levers		GO, SBE	S/N	3210573-0048, 3216020-0175, shiny12, dull24, none

Frame w/ s/s cage							12 pl SBE alum old/new, 12 pl s/s, 24pl SBE alum, bk1, bk2

Niskins		# on frame: 12 	      1.5, 5, 10, 30 L	Go-Flo, Ext resist OTE / GO, Internal resist, w/ Hiflow Valves, none

Winch / wire						#	1in, 1out, 2, 3, Scanfish / .322”, .500”, .680”, 3/8”, ½”, 9/16”

Slip Rings (Meridian, Focal)				S/N	6.859, 8.538 (10wire), 8.594, Focal 1801308, none

Pinger	(Williams)					S/N	W1, W2, W3, none

(Benthos 2216, 12000m)				S/N	1237, none

Note for SBE 9plus CTD underwater unit S/N 12293-0444, 0523 & 0607:

CTDs 0444, 0523, and 0607 were supplied with an isolated 12-volt power source for a Sea Tech fluorometer in a 3000-meter stainless steel housing. When this fluorometer is used it must be installed on connector JT2 (see drawing 50077). This will put the fluorometer output on A/D channel 1 (External Voltage 1).

Underwater unit S/N 9P16940-0607, as of x x 2000:

Pressure sensor range						0-10,000 psia, 0-6885 d-Bar

pressure sensor, Digiquartz (with temp comp, SBE)		S/N	80016

Modulo							S/N	MOD12P-439

AD590M (enter in Seacon)					0.012837, , .0128452

AD590B (enter in Seacon)						-9.963864, -8.66958

A/D Input Voltage Range						0 to 5 volts DC

Logic Board EPROM				Version 1.0

Modem Interface NOT installed

NO Modem Board Microcontroller			Version 2.0A, Version 1.1		P/N N/A, 11124, 11194

Isolated Fluorometer Power Interface NOT Installed

## Voltages in Deck test437xapooFs1_0607.dat		Open		Blocked		Capped

External voltage 0	Transmissometer 480DR		4.8083		0.0574

External voltage 1	spare				0.000

External voltage 2	Alt 1075				4.926

External voltage 3	PAR 4479			~4.1612daylight	0.4615		~1.9389 w/hole in top

External voltage 4	DO 1230			1.355

External voltage 5	DO 0497			1.402

External voltage 6	spare				4.923

External voltage 7	Fluor 231			0.0611		4.9145w/fluor stick  1.538

External voltage 8	spare

External voltage 9	SPAR 20190			4.497daylight	4.994